## Evidence
Row | ID | Section | Claim | Verdict | Finding | Verbatim passage | Passage location | Design / N | Source | Reference | PMID | DOI | Search | Verified | Notes | Source status
R001 | C01 | Cuaderno Cardiología/IABP | Errores de temporización del IABP (inflado precoz, inflado tardío, desinflado precoz, desinflado tardío): cómo se ven en la curva de presión arterial y qué consecuencias hemodinámicas tienen | relevant | Inflado precoz: la válvula aórtica se cierra antes de tiempo, sube la poscarga, queda más volumen telesistólico, baja el GC; en la curva, el aumento diastólico cae antes de la muesca dicrótica y pierde la «V» | Early inflation will cause the IABP to inflate before the AV has closed. This has the potential to cause the AV to close prematurely, because of increased afterload, leading to impaired LV emptying and possibly aortic regurgitation. Incomplete LV emptying will lead to an increase in LV end-diastolic volume, LVEDP, and pulmonary capillary wedge pressure, which will increase myocardial oxygen demand. Early inflation will lead to diastolic augmentation occurring prior to the dicrotic notch with loss of the desired “V” shape. The diastolic augmentation may encroach onto the unassisted systolic peak to such a degree that the 2 peaks become difficult to differentiate. | Sección Early inflation | Revisión narrativa (JSCAI 2024) | secondary | Gillespie LE, Lane BH, Shaw CR, Gorder K, Grisoli A, Lavallee M, et al. The Intra-aortic Balloon Pump: A Focused Review of Physiology, Transport Logistics, Mechanics, and Complications. J Soc Cardiovasc Angiogr Interv. 2024;3(5):101337. doi:10.1016/j.jscai.2024.101337 | 39132456 | 10.1016/j.jscai.2024.101337 | PubMed 2026-10-27: (intra-aortic balloon[tiab] OR intraaortic balloon[tiab] OR IABP[tiab]) AND timing[tiab] AND (inflation[tiab] OR deflation[tiab] OR waveform*[tiab]) = 48; ampliada con counterpulsation/troubleshooting/review[pt] = 131 (poco específica); (...) AND (early/late/premature inflation/deflation, delayed deflation) = 24; PMC OA full-text search aportó PMID 39132456 y 36263076 |  |  | fetched
R002 | C01 | Cuaderno Cardiología/IABP | Errores de temporización del IABP (inflado precoz, inflado tardío, desinflado precoz, desinflado tardío): cómo se ven en la curva de presión arterial y qué consecuencias hemodinámicas tienen | relevant | Inflado tardío: aumento diastólico subóptimo con perfusión coronaria inadecuada; en la curva, el inflado ocurre bien después de la muesca dicrótica y se pierde la «V» | Late inflation of the IABP occurs well after closure of the AV. This has the potential to cause inadequate perfusion of the coronary arteries due to suboptimal diastolic augmentation. The arterial waveform will show IABP inflation occurring well after the dicrotic notch, again resulting in loss of the optimal V shape. | Sección Late inflation | Revisión narrativa (JSCAI 2024) | secondary | Gillespie LE, Lane BH, Shaw CR, Gorder K, Grisoli A, Lavallee M, et al. The Intra-aortic Balloon Pump: A Focused Review of Physiology, Transport Logistics, Mechanics, and Complications. J Soc Cardiovasc Angiogr Interv. 2024;3(5):101337. doi:10.1016/j.jscai.2024.101337 | 39132456 | 10.1016/j.jscai.2024.101337 | PubMed 2026-10-27: (intra-aortic balloon[tiab] OR intraaortic balloon[tiab] OR IABP[tiab]) AND timing[tiab] AND (inflation[tiab] OR deflation[tiab] OR waveform*[tiab]) = 48; ampliada con counterpulsation/troubleshooting/review[pt] = 131 (poco específica); (...) AND (early/late/premature inflation/deflation, delayed deflation) = 24; PMC OA full-text search aportó PMID 39132456 y 36263076 |  |  | fetched
R003 | C01 | Cuaderno Cardiología/IABP | Errores de temporización del IABP (inflado precoz, inflado tardío, desinflado precoz, desinflado tardío): cómo se ven en la curva de presión arterial y qué consecuencias hemodinámicas tienen | relevant | Desinflado precoz: menor reducción de poscarga y de demanda de O2; posible flujo coronario/carotídeo retrógrado; en la curva, caída brusca de presión justo tras el pico de aumento diastólico y PAS asistida >= no asistida | Early deflation of IABP during the diastolic phase ends before isovolumetric contraction occurs. This decreases the efficacy of the IABP to reduce afterload, reducing the boost in CPP and obviating the potential reduction in myocardial oxygen demand. Early deflation may lead to retrograde coronary and carotid blood flow due to the vacuum-like effect in mid-diastole. The arterial waveform will show a sharp drop in pressure soon after the peak of diastolic augmentation. The assisted SBP may be equal to or higher than the unassisted SBP and the assisted EDBP may be equal to or less than the unassisted EDBP. | Sección Early deflation | Revisión narrativa (JSCAI 2024) | secondary | Gillespie LE, Lane BH, Shaw CR, Gorder K, Grisoli A, Lavallee M, et al. The Intra-aortic Balloon Pump: A Focused Review of Physiology, Transport Logistics, Mechanics, and Complications. J Soc Cardiovasc Angiogr Interv. 2024;3(5):101337. doi:10.1016/j.jscai.2024.101337 | 39132456 | 10.1016/j.jscai.2024.101337 | PubMed 2026-10-27: (intra-aortic balloon[tiab] OR intraaortic balloon[tiab] OR IABP[tiab]) AND timing[tiab] AND (inflation[tiab] OR deflation[tiab] OR waveform*[tiab]) = 48; ampliada con counterpulsation/troubleshooting/review[pt] = 131 (poco específica); (...) AND (early/late/premature inflation/deflation, delayed deflation) = 24; PMC OA full-text search aportó PMID 39132456 y 36263076 |  |  | fetched
R004 | C01 | Cuaderno Cardiología/IABP | Errores de temporización del IABP (inflado precoz, inflado tardío, desinflado precoz, desinflado tardío): cómo se ven en la curva de presión arterial y qué consecuencias hemodinámicas tienen | relevant | Desinflado tardío: balón aún inflado tras abrirse la válvula; aumenta la poscarga, prolonga la contracción isovolumétrica y la demanda de O2; en la curva, PTD asistida >= no asistida y pico diastólico ensanchado | Deflation should end near the end of diastole, during isovolumetric contraction; if IABP deflation is prolonged, it may still be in the process of deflating after the AV has opened. In this scenario, rather than reducing afterload, the resistance from a still-inflated balloon during systole will increase afterload, interfering with LV ejection, and prolonging isovolumetric contraction. These effects will increase myocardial oxygen demand. The characteristic arterial waveform of late deflation will show an assisted end-diastolic pressure that is equal to or greater than the unassisted end-diastolic pressure (rather than appropriately decreased) and a widened augmented diastolic peak. | Sección Late deflation | Revisión narrativa (JSCAI 2024) | secondary | Gillespie LE, Lane BH, Shaw CR, Gorder K, Grisoli A, Lavallee M, et al. The Intra-aortic Balloon Pump: A Focused Review of Physiology, Transport Logistics, Mechanics, and Complications. J Soc Cardiovasc Angiogr Interv. 2024;3(5):101337. doi:10.1016/j.jscai.2024.101337 | 39132456 | 10.1016/j.jscai.2024.101337 | PubMed 2026-10-27: (intra-aortic balloon[tiab] OR intraaortic balloon[tiab] OR IABP[tiab]) AND timing[tiab] AND (inflation[tiab] OR deflation[tiab] OR waveform*[tiab]) = 48; ampliada con counterpulsation/troubleshooting/review[pt] = 131 (poco específica); (...) AND (early/late/premature inflation/deflation, delayed deflation) = 24; PMC OA full-text search aportó PMID 39132456 y 36263076 |  |  | fetched
R005 | C01 | Cuaderno Cardiología/IABP | Errores de temporización del IABP (inflado precoz, inflado tardío, desinflado precoz, desinflado tardío): cómo se ven en la curva de presión arterial y qué consecuencias hemodinámicas tienen | relevant | Resumen de los cuatro errores: inflado precoz o desinflado tardío aumentan poscarga y demanda de O2; inflado tardío o desinflado precoz no aumentan lo suficiente la perfusión coronaria (sin descripción de la curva) | Early inflation or late deflation leads to increased afterload and increased myocardial oxygen demand. Late inflation or early deflation does not sufficiently increase coronary perfusion. | Texto completo (sección de IABP) | Revisión narrativa (Int J Heart Fail 2020) | secondary | Choi MS, Shim H, Cho YH. Mechanical Circulatory Support for Acute Heart Failure Complicated by Cardiogenic Shock. Int J Heart Fail. 2020;2(1):23-44. doi:10.36628/ijhf.2019.0015 | 36263076 | 10.36628/ijhf.2019.0015 | PubMed 2026-10-27: (intra-aortic balloon[tiab] OR intraaortic balloon[tiab] OR IABP[tiab]) AND timing[tiab] AND (inflation[tiab] OR deflation[tiab] OR waveform*[tiab]) = 48; ampliada con counterpulsation/troubleshooting/review[pt] = 131 (poco específica); (...) AND (early/late/premature inflation/deflation, delayed deflation) = 24; PMC OA full-text search aportó PMID 39132456 y 36263076 |  |  | fetched
R006 | C01 | Cuaderno Cardiología/IABP | Errores de temporización del IABP (inflado precoz, inflado tardío, desinflado precoz, desinflado tardío): cómo se ven en la curva de presión arterial y qué consecuencias hemodinámicas tienen | relevant | Matiz sobre el desinflado tardío: retrasarlo al inicio del QRS aumentó el flujo cerebral un 20 % y el GC un 15 % sin prolongar la contracción isovolumétrica (la «sensación de vacío» del desinflado causa inversión de flujo carotídeo) | Delaying IABP deflation trigger to the beginning of QRS effectively increased the cerebral blood flow by 20%, mostly by reducing the reverse component flow caused by the diastolic balloon deflation. Extending the deflation to the early systole was safe and favourably impacted on cardiac mechanics, increasing CO by 15% without prolonging LV isovolumetric contraction and ejection phases. | Abstract | Estudio clínico hemodinámico | primary | Boccellino A, Sacchi S, Pazzanese V, Calvo F, Gramegna M, Ingallina G, et al. Effects of Intra-Aortic Balloon Pump delayed deflation timing on carotid blood flow and cardiac mechanics: An ultrasound and haemodynamic study. J Heart Lung Transplant. 2023;42(4):451-455. doi:10.1016/j.healun.2022.12.007 | 36682892 | 10.1016/j.healun.2022.12.007 | PubMed 2026-10-27: (intra-aortic balloon[tiab] OR intraaortic balloon[tiab] OR IABP[tiab]) AND timing[tiab] AND (inflation[tiab] OR deflation[tiab] OR waveform*[tiab]) = 48; ampliada con counterpulsation/troubleshooting/review[pt] = 131 (poco específica); (...) AND (early/late/premature inflation/deflation, delayed deflation) = 24; PMC OA full-text search aportó PMID 39132456 y 36263076 |  |  | fetched
R007 | C01 | Cuaderno Cardiología/IABP | Errores de temporización del IABP (inflado precoz, inflado tardío, desinflado precoz, desinflado tardío): cómo se ven en la curva de presión arterial y qué consecuencias hemodinámicas tienen | relevant | La curva de la consola llega con retraso y deformada (inflado ~74 ms y desinflado ~71 ms frente a la raíz aórtica; ~170 ms frente a la radial; 11 pacientes) | Delay and distortion of the arterial waveform was consistently found when measured through the IABP catheter lumen. | abstract | Estudio clínico, n=11 | primary | Pantalos GM, Koenig SC, Gillars KJ, Haugh GS, Dowling RD, Gray LA Jr. Intraaortic balloon pump timing discrepancies in adult patients. Artif Organs. 2011;35(9):857-66. doi:10.1111/j.1525-1594.2011.01292.x | 21883317 | 10.1111/j.1525-1594.2011.01292.x | PubMed 2026-10-27: (intra-aortic balloon[tiab] OR intraaortic balloon[tiab] OR IABP[tiab]) AND timing[tiab] AND (inflation[tiab] OR deflation[tiab] OR waveform*[tiab]) = 48; ampliada con counterpulsation/troubleshooting/review[pt] = 131 (poco específica); (...) AND (early/late/premature inflation/deflation, delayed deflation) = 24; PMC OA full-text search aportó PMID 39132456 y 36263076 |  |  | fetched

## Seen, not added
Row | ID | Section | Claim | Verdict | Finding | Verbatim passage | Passage location | Design / N | Source | Reference | PMID | DOI | Search | Verified | Notes | Source status

## _sources (hidden)
key | chunk | pmid | pmcid | doi | status | meta | text
PMID:16881490 | 0 | 16881490 |  | 10.1016/s0001-2092(06)60096-3 | fetched | {"title": "Intra-aortic balloon pump therapy--a primer for perioperative nurses.", "authors": ["Tremper RS"], "journal_abbrev": "AORN J", "year": 2006, "volume": "84", "issue": "1", "pages": "34-44; quiz 45-8", "doi": "10.1016/s0001-2092(06)60096-3", "pmid": "16881490", "pmcid": null, "pub_types": ["Journal Article", "Review"]} | Intra-aortic balloon pump therapy--a primer for perioperative nurses. The intra-aortic balloon pump (IABP) is a commonly used circulatory-assist device that increases myocardial oxygen supply and reduces myocardial oxygen demand, thus improving left ventricular function. The inflation and deflation pattern of the IABP is the opposite of the heart's rhythm. As the heart rests, the balloon inflates, displacing blood volume into the coronary sinus and systemic circulation. The balloon deflates just before ventricular ejection, decreasing the pressure that the ventricle must eject against. This article describes indications and contraindications for use of the IABP, insertion and timing of the device, and troubleshooting measures for common malfunctions.
PMID:8194350 | 0 | 8194350 |  |  | fetched | {"title": "Assessing the five points of the intra-aortic balloon pump waveform.", "authors": ["Wojner AW"], "journal_abbrev": "Crit Care Nurse", "year": 1994, "volume": "14", "issue": "3", "pages": "48-52", "doi": null, "pmid": "8194350", "pmcid": null, "pub_types": ["Journal Article"]} | Assessing the five points of the intra-aortic balloon pump waveform. IABP timing challenges the nurse not only with a waveform analysis but also an understanding of complex physiologic concepts crucial to patient management. Optimizing timing through assessment of the five points of the IABP waveform permits the critical care nurse to increase the hemodynamic benefits of IABP therapy, potentially improving patient outcome.
PMID:9095807 | 0 | 9095807 |  |  | fetched | {"title": "Timing waveform analysis.", "authors": ["Joseph DL", "Spadoni SM"], "journal_abbrev": "Crit Care Nurs Clin North Am", "year": 1996, "volume": "8", "issue": "4", "pages": "349-56", "doi": null, "pmid": "9095807", "pmcid": null, "pub_types": ["Journal Article", "Review"]} | Timing waveform analysis. This article provides a description of correct balloon pump timing as well as the potential timing errors that can occur. There are many factors in addition to timing that can affect the altered arterial waveform during balloon pumping. When analyzing the waveform for appropriate timing, it is important to consider factors associated with the patient, the IAB, and the IABP console. Patient treatment decisions should be based on the complete clinical picture, not solely on the characteristics of the altered arterial waveform.
PMID:8811147 | 0 | 8811147 |  |  | fetched | {"title": "Intra-aortic balloon counterpulsation timing.", "authors": ["Cadwell CA", "Quaal SJ"], "journal_abbrev": "Am J Crit Care", "year": 1996, "volume": "5", "issue": "4", "pages": "254-61; quiz 262-3", "doi": null, "pmid": "8811147", "pmcid": null, "pub_types": ["Journal Article"]} | Intra-aortic balloon counterpulsation timing. Intra-aortic balloon counterpulsation is a product of an endeavor begun more than 3 decades ago. First the concept, then applications of inflating a nonthrombogenic balloon within the aorta during diastole were explored. The intra-aortic balloon is timed to inflate with aortic valve closure. Two options are available for deflation timing: (1) conventional timing, where deflation is estimated to be completed at some point during isovolumetric contraction, prior to ejection; and (2) R wave deflation, also known as real timing, where deflation is triggered by each QRS complex to occur during isovolumetric contraction or early systole. Although clinical implementation of both conventional and real timing was introduced in 1968, limited information about these methods of deflation is currently available. This article elucidates the differences between these models and suggests implications for clinical practice and further research.
PMID:10148160 | 0 | 10148160 |  |  | fetched | {"title": "Techniques for IABP timing, use--and discontinuance. Counterpulsation can reduce ischemia and improve hemodynamics.", "authors": ["Sorrentino M", "Feldman T"], "journal_abbrev": "J Crit Illn", "year": 1992, "volume": "7", "issue": "4", "pages": "597-604", "doi": null, "pmid": "10148160", "pmcid": null, "pub_types": ["Journal Article"]} | Techniques for IABP timing, use--and discontinuance. Counterpulsation can reduce ischemia and improve hemodynamics. By restoring the balance between myocardial oxygen supply and demand, intra-aortic balloon counterpulsation can aid the ischemic or failing left ventricle. Rapid inflation and deflation of the balloon reduces afterload and improves coronary perfusion. Timing of inflation and deflation is adjusted according to the arterial pressure waveform. The balloon should inflate just after the aortic valve closes and deflate just before the aortic valve opens; early or late balloon inflation or deflation blunts the therapeutic effects. Discontinue the IABP support gradually; monitor the patient frequently for any signs of hemodynamic compromise, redevelopment of ischemic chest pain, or electrocardiographic changes.
PMID:2818886 | 0 | 2818886 |  |  | fetched | {"title": "Perspectives on intra-aortic balloon-pump timing.", "authors": ["Gould KA"], "journal_abbrev": "Crit Care Nurs Clin North Am", "year": 1989, "volume": "1", "issue": "3", "pages": "469-73", "doi": null, "pmid": "2818886", "pmcid": null, "pub_types": ["Journal Article"]} | Perspectives on intra-aortic balloon-pump timing. Triggering and timing of the intra-aortic balloon pump are becoming more sophisticated but more clinically manageable. The key to understanding the data currently available is to go beyond early instruction to "eyeball" waveform analysis. Conclusive assessment of the patient's response to therapy lies in a greater understanding of today's technology and clinical competency in evaluating the patient's total hemodynamic, clinical, and laboratory response to assure maximal therapeutic efficiency of counterpulsation therapy. Nursing research on an ongoing basis will contribute to a greater understanding of this complex therapy. Many of the principles of balloon pumping were derived in the years following its initial introduction. Since that time, we have gained a depth of knowledge and an increased capacity for measuring physiologic and hemodynamic responses at the bedside. Concurrently, the medical industry has designed faster, smarter, and more efficient equipment to complement patient care. The critical care nurse faces the challenge of synthesizing the newest information from both areas and developing a greater understanding of the ongoing care of the patient on the IABP.
PMID:4003877 | 0 | 4003877 |  | 10.1007/BF02584236 | fetched | {"title": "Control of intraaortic balloon pumping: theory and guidelines for clinical applications.", "authors": ["Jaron D", "Moore TW", "He P"], "journal_abbrev": "Ann Biomed Eng", "year": 1985, "volume": "13", "issue": "2", "pages": "155-75", "doi": "10.1007/BF02584236", "pmid": "4003877", "pmcid": null, "pub_types": ["Journal Article", "Research Support, U.S. Gov't, Non-P.H.S.", "Research Support, U.S. Gov't, P.H.S."]} | Control of intraaortic balloon pumping: theory and guidelines for clinical applications. The effectiveness of intraaortic balloon pumping was investigated by using a lumped parameter model of the cardiovascular/assist device system. The model consists of a time-varying elastance left ventricular simulation, a 2-element windkessel arterial simulation, and an RC venous return and pulmonary simulation. The four major hemodynamic variables, stroke volume (SV), aortic mean diastolic pressure (MDP), tension time index (TTI), and aortic end diastolic pressure (EDP), were divided into two categories related to system energy supply and demand: "external" and "internal" variables. The effects of balloon pumping on these variables can be described by closed-form equations that yield an optimal solution. The model prediction suggests that, in the ideal case, optimization of balloon pumping calls for instantaneous inflation of the balloon to maximum volume at end systole and instantaneous complete deflation at end diastole. For finite inflation/deflation rates, the optimal time for the start of inflation is end systole. Deflation timing, however, involves a tradeoff between maximizing the external variables and minimizing the internal variables. These predictions were tested using a nonlinear digital computer model. The results also suggest that when SV is not being monitored, optimal inflation timing can be controlled from the measurements of TTI or pulmonary venous pressure; optimal deflation timing can be controlled by a weighted combination of MDP and EDP.
PMID:17672189 | 0 | 17672189 | PMC4680672 |  | fetched | {"title": "Evaluating the efficacy of intra-aortic balloon pump timing using the auto-timing mode of operation with the Datascope CS100.", "authors": ["Osentowski MK", "Holt DW"], "journal_abbrev": "J Extra Corpor Technol", "year": 2007, "volume": "39", "issue": "2", "pages": "87-90", "doi": null, "pmid": "17672189", "pmcid": "PMC4680672", "pub_types": ["Evaluation Study", "Journal Article"]} | Evaluating the efficacy of intra-aortic balloon pump timing using the auto-timing mode of operation with the Datascope CS100. The intra-aortic balloon pump (IABP) was first introduced by Moulopoulos et al. in 1962 and was described clinically by Kantrowitz et al. in 1968. The subsequent development of the IABP has resulted in its widespread application, resulting in >70,000 applications per year, making it the most common cardiac assist device after pharmacologic intervention. The correct timing of the inflation and deflation of this device is critical to it being clinically advantageous to the patient. The purpose of this study was primarily to examine the Datascope CS100 automated timing of inflation and deflation, and secondarily, to create an awareness of the risks of solely relying on technology for patient treatment decisions in lieu of clinical evaluative skills. Timing data were collected from 165 IABP patients from January 1, 2003 to March 31, 2005 using the Datascope series System 97, System 98, and CS100 IABP consoles. Timing criteria used in the evaluation of the IABP in this study included proper inflation, defined as inflation at the dicrotic notch of arterial pressure waveform, and proper deflation, resulting in a presystolic decrease of diastolic pressure by 5-10 mmHg. These parameters were evaluated by examining a printed strip from the IABP console after timing selection stabilization. Before implementation of the CS100, 78.6% of inflations and 53.6% of deflations were correct when using the criteria listed above. After implementation of the CS100, 83.3% of inflations were correct and 44.4% of deflations were correct. Although improvement in consistency of IABP inflation quality was observed, a net decrease in the quality of IABP deflations was also observed. Total reliance on the autotiming operation mode of the CS100 may not be warranted, and imposing clinical judgment in therapeutic application is necessary to avoid dangerous timing errors. An awareness of the dangers of solely relying on technology for patient treatment decisions vs. the use of technology combined with clinical evaluative skills for assessing therapeutic intervention to improve patient care is discussed.
PMID:21883317 | 0 | 21883317 |  | 10.1111/j.1525-1594.2011.01292.x | fetched | {"title": "Intraaortic balloon pump timing discrepancies in adult patients.", "authors": ["Pantalos GM", "Koenig SC", "Gillars KJ", "Haugh GS", "Dowling RD", "Gray LA Jr"], "journal_abbrev": "Artif Organs", "year": 2011, "volume": "35", "issue": "9", "pages": "857-66", "doi": "10.1111/j.1525-1594.2011.01292.x", "pmid": "21883317", "pmcid": null, "pub_types": ["Journal Article", "Research Support, Non-U.S. Gov't"]} | Intraaortic balloon pump timing discrepancies in adult patients. The objective of this clinical study was to quantify the incidence and magnitude of intraaortic balloon pump (IABP) inflation and deflation landmark discrepancies associated with the IABP catheter arterial pressure waveform. Cardiac surgery patients with an IABP inserted prior to surgery were recruited. Following cardiac exposure, a high-fidelity pressure catheter was inserted into the aortic root for digital recording. The radial artery pressure signal was simultaneously recorded from the patient monitor along with the arterial pressure and electrocardiogram waveforms from the IABP console while operating at 1:1 and 1:2 synchronization. In selected patients, recordings were obtained with the IABP timed to the high-fidelity aortic root waveform. In all 11 patients, inflation and deflation landmark delays were observed when comparing the aortic root waveforms to the IABP arterial pressure waveforms (inflation delay = 74 +/- 29 [23-117] ms; deflation delay = 71 +/- 37 [24-141] ms, mean +/- standard deviation [min-max]). Delays were greater when compared to the radial artery waveform (inflation delay = 175 +/- 50 [100-233] ms; deflation delay = 168 +/- 52 [100-274] ms). In all cases, the landmark delays were statistically different from zero (P < 0.001). Diastolic augmentation and afterload reduction varied with waveform source. Conflicting indications of afterload reduction occurred in four patients. Timing to the aortic root waveform resulted in greater diastolic pressure augmentation and afterload reduction but mixed changes in stroke volume. Delay and distortion of the arterial waveform was consistently found when measured through the IABP catheter lumen. These delays can alter IABP efficacy and may be eliminated by using high-fidelity sensing of aortic pressure.
PMID:25744191 | 0 | 25744191 |  | 10.5301/ijao.5000385 | fetched | {"title": "Does conventional intra-aortic balloon pump trigger timing produce optimal hemodynamic effects in vivo?", "authors": ["Kolyva C", "Pantalos GM", "Pepper JR", "Khir AW"], "journal_abbrev": "Int J Artif Organs", "year": 2015, "volume": "38", "issue": "3", "pages": "146-53", "doi": "10.5301/ijao.5000385", "pmid": "25744191", "pmcid": null, "pub_types": ["Journal Article", "Research Support, Non-U.S. Gov't"]} | Does conventional intra-aortic balloon pump trigger timing produce optimal hemodynamic effects in vivo? PURPOSE: The intra-aortic balloon pump (IABP) provides circulatory support through counterpulsation. The hemodynamic effects of the IABP may vary with assisting frequency and depend on IAB inflation/deflation timing. We aimed to assess in vivo the IABP benefits on coronary, aortic, and left ventricular hemodynamics at different assistance frequencies and trigger timings. METHODS: Six healthy, anesthetized, open-chest sheep received IABP support at 5 timing modes (EC, LC, CC, CE, CL, corresponding to early/late/conventional/conventional/conventional inflation and conventional/conventional/conventional/early/late deflation, respectively) with frequency 1:3 and 1:1. Aortic (Q(ao)) and coronary (Q(cor)) flow, and aortic (P(ao)) and left ventricular (PLV) pressure were recorded simultaneously, with and without IABP support. Integrating systolic Q(ao) yielded stroke volume (SV). RESULTS: EC at 1:1 produced the lowest end-diastolic P(ao) (59.5 +/- 7.8 mmHg [EC], 63.4 +/- 11.1 mmHg [CC]), CC at 1:1 the lowest systolic PLV (69.1 +/- 6.5 mmHg [CC], 76.4 +/- 6.5 mmHg [control]), CC at 1:1 the highest SV (88.5 +/- 34.4 ml [CC], 76.6 +/- 31.9 ml [control]) and CC at 1:3 the highest diastolic Qcor (187.2 +/- 25.0 ml/min [CC], 149.9 +/- 16.6 ml/min [control]). Diastolic P(ao) augmentation was enhanced by both assistance frequencies alike, and optimal timings were EC for 1:3 (10.4 +/- 2.8 mmHg [EC], 6.7 +/- 3.8 mmHg [CC]) and CC for 1:1 (10.8 +/- 6.7 mmHg [CC], -3.0 +/- 3.8 mmHg [control]). CONCLUSIONS: In our experiments, neither a single frequency nor a single inflation/deflation timing, including conventional IAB timing, has shown superiority by uniformly benefiting all studied hemodynamic parameters. A choice of optimal frequency and IAB timing might need to be made based on individual patient hemodynamic needs rather than as a generalized protocol.
PMID:41096015 | 0 | 41096015 | PMC12524590 | 10.3390/jcm14196935 | fetched | {"title": "Cardiovascular Physiology During Mechanical Circulatory Support: Implications for Management and Monitoring.", "authors": ["Crimi E", "Rajkumar K", "Coleman S", "Fernando R", "Marchant B", "Garner C", "Gaillard J", "Hicks MH", "Maves RC", "Khanna AK"], "journal_abbrev": "J Clin Med", "year": 2025, "volume": "14", "issue": "19", "pages": null, "doi": "10.3390/jcm14196935", "pmid": "41096015", "pmcid": "PMC12524590", "pub_types": ["Journal Article", "Review"]} | Cardiovascular Physiology During Mechanical Circulatory Support: Implications for Management and Monitoring. Background/Objectives: Mechanical circulatory support (MCS) is increasingly utilized for the management of acute decompensated heart failure (HF) and cardiogenic shock (CS). The primary goals of MCS are to restore systemic perfusion, reduce cardiac workload, and support end-organ function. A thorough understanding of cardiovascular physiology in patients supported by MCS is essential for clinical decision-making. This review summarizes current evidence on the physiological effects of various MCS devices, key monitoring techniques, patient management, and explores the emerging role of artificial intelligence (AI) in this field. Main Text: Short-term MCS devices include intra-aortic balloon pumps (IABP), percutaneous left-sided devices such as Impella (Abiomed, Danvers, MA, USA) and TandemHeart (LivaNova, London, UK), percutaneous right-sided support devices like Protek Duo (LivaNova, London, UK) and Impella RP Flex (Abiomed, Danvers, MA, USA), and veno-arterial extracorporeal membrane oxygenation (VA-ECMO). Long-term support is mainly provided by left ventricular assist devices (LVADs), including the HeartMate 3 (Abbott Laboratories, Chicago, IL, USA). Optimal MCS application requires an understanding of device-specific cardiovascular interactions and expertise in appropriate monitoring tools to assess device performance and patient response. The choice of device, timing of initiation, and patient selection must be individualized, with careful consideration of ethical implications. The integration of AI offers significant potential to advance clinical care by improving complication prediction, enabling real-time optimization of device settings, and refining patient selection criteria. Conclusions: MCS is a rapidly evolving field that requires a comprehensive understanding of cardiovascular interactions, careful selection of monitoring strategies, and individualized clinical management. Future research should address current device limitations, clarify device-specific clinical applications, and assess the validity of AI-driven technologies. ABSTRACT: Background/Objectives: Mechanical circulatory support (MCS) is increasingly utilized for the management of acute decompensated heart failure (HF) and cardiogenic shock (CS). The primary goals of MCS are to restore systemic perfusion, reduce cardiac workload, and support end-organ function. A thorough understanding of cardiovascular physiology in patients supported by MCS is essential for clinical decision-making. This review summarizes current evidence on the physiological effects of various MCS devices, key monitoring techniques, patient management, and explores the emerging role of artificial intelligence (AI) in this field. Main Text: Short-term MCS devices include intra-aortic balloon pumps (IABP), percutaneous left-sided devices such as Impella (Abiomed, Danvers, MA, USA) and TandemHeart (LivaNova, London, UK), percutaneous right-sided support devices like Protek Duo (LivaNova, London, UK) and Impella RP Flex (Abiomed, Danvers, MA, USA), and veno-arterial extracorporeal membrane oxygenation (VA-ECMO). Long-term support is mainly provided by left ventricular assist devices (LVADs), including the HeartMate 3 (Abbott Laboratories, Chicago, IL, USA). Optimal MCS application requires an understanding of device-specific cardiovascular interactions and expertise in appropriate monitoring tools to assess device performance and patient response. The choice of device, timing of initiation, and patient selection must be individualized, with careful consideration of ethical implications. The integration of AI offers significant potential to advance clinical care by improving complication prediction, enabling real-time optimization of device settings, and refining patient selection criteria. Conclusions: MCS is a rapidly evolving field that requires a comprehensive understanding of cardiovascular interactions, careful selection of monitoring strategies, and individualized clinical management. Future research should address current device limitations, clarify device-specific clinical applications, and assess the validity of AI-driven technologies. 1. INTRODUCTION: Mechanical circulatory support (MCS) is increasingly used in the multidisciplinary management of decompensated heart failure (HF) and cardiogenic shock (CS). These complex clinical syndromes are characterized by impaired end-organ perfusion and high mortality rates ranging from 30% to 50% [1,2,3,4]. The primary goals of MCS are to restore systemic perfusion, decrease cardiac workload and myocardial oxygen demand, support end-organ function, and minimize prolonged high-dose inotropic support, which may exacerbate organ dysfunction [5]. Applications of MCS devices span from acute short-term stabilization to longer-term support aimed at bridging patients to myocardial recovery, transplantation, or destination therapy. Timely and individualized implementation of MCS strategies is critical to prevent progression of end-organ dysfunction and improve clinical outcomes [6,7,8,9]. A thorough understanding of cardiovascular physiology associated with MCS devices is essential for optimizing clinical management, patient selection, and monitoring strategies. This review describes the distinct physiological effects of each MCS device, highlights key invasive and non-invasive monitoring techniques, examines clinical management and ethical considerations, and explores the role of artificial intelligence in optimizing MCS strategies. 2.1. DEFINITIONS OF HEART FAILURE AND CARDIOGENIC SHOCK: HF is a clinical syndrome resulting from structural and/or functional cardiac abnormalities that impair ventricular filling or ejection. HF is commonly categorized as HF with reduced ejection fraction (HFrEF), defined as left ventricular ejection fraction (LVEF) ≤40%; HF with mildly reduced ejection fraction (HFmrEF), defined as LVEF 41–49% and sharing characteristics of both reduced and preserved phenotypes; and HF with preserved ejection fraction (HFpEF), defined as LVEF ≥50% with objective evidence of increased left ventricular filling pressures such as elevated natriuretic peptides or abnormal hemodynamic indices [1]. The ACC/AHA stages of HF outline disease progression. Stage A identifies patients at risk for HF but without symptoms or structural abnormalities. Stage B (Pre-HF) refers to those with asymptomatic structural changes, abnormal biomarkers, or elevated filling pressures. Stage C is defined by structural heart disease with current or prior symptoms of HF. Stage D (advanced or end-stage HF) includes severe, refractory symptoms that interfere with daily life, recurrent hospitalizations, and poor prognosis despite optimal medical and device-based therapy [1]. End-stage HF (Stage D) represents advanced heart failure that is refractory to standard, guideline-directed medical and device-based therapies [10]. Patients in this category frequently require advanced interventions such as durable mechanical circulatory support (MCS) or cardiac transplantation. To further stratify advanced HF, the INTERMACS profiles classify patients according to severity and urgency, ranging from critical cardiogenic shock (Profile 1) to advanced but stable disease (Profile 7) as shown in Table 1 [1,11]. Acute decompensated HF (ADHF) refers to an acute or subacute worsening of signs and symptoms that require urgent medical therapy or hospitalization. Clinical presentation is usually dominated by pulmonary and systemic congestion, often in the setting of preserved or elevated blood pressure. Perfusion is generally maintained, which distinguishes ADHF from cardiogenic shock [1,11]. CS is the most severe manifestation of acute HF, defined as inadequate cardiac output resulting in systemic hypoperfusion. Although hypotension (systolic blood pressure <90 mmHg or vasopressor dependence) is a hallmark, diagnosis also requires evidence of end-organ compromise such as oliguria, altered mental status, metabolic acidosis, or hepatic and renal dysfunction [4]. Acute myocardial infarction remains the most common precipitating cause of cardiogenic shock, though myocarditis, decompensated chronic heart failure, and postcardiotomy failure are also important etiologies [10]. The underlying pathophysiology may involve decreased contractility (e.g., ischemic cardiomyopathy or myocarditis), increased afterload, or inadequate LV preload due to RV dysfunction. The Society for Cardiovascular Angiography and Interventions (SCAI) shock stages provide a standardized framework for classification, ranging from patients at risk (Stage A) to those with profound circulatory collapse (Stage E) as shown in Table 2 [2]. 2.2. NEUROHORMONAL MECHANISMS IN HEART FAILURE AND CARDIOGENIC SHOCK: Heart failure and cardiogenic shock are driven by maladaptive neurohormonal activation that promotes adverse remodeling and multi-organ dysfunction. The most important systems are the sympathetic nervous system (SNS), the renin–angiotensin–aldosterone system (RAAS), and arginine vasopressin (AVP), partially counterbalanced by natriuretic peptides, nitric oxide, and other vasodilatory mediators [12,13]. SNS activation occurs early through impaired baroreflex control. Circulating norepinephrine levels increase with severity and predict mortality [14]. While adrenergic drive acutely supports contractility and perfusion, sustained stimulation depletes myocardial norepinephrine, downregulates β-receptors, and disrupts calcium cycling, leading to excess oxygen demand, arrhythmias, and fibrosis [12]. Renal hypoperfusion and adrenergic tone trigger RAAS activation. Angiotensin II and aldosterone initially maintain arterial pressure through vasoconstriction and sodium retention. Chronically, they promote hypertrophy, fibrosis, oxidative stress, and endothelial dysfunction. The ACE2/Ang-(1–7)/Mas receptor axis provides vasodilatory and antifibrotic counterbalance but is insufficient in advanced disease. AVP levels are often inappropriately high in HF, causing V1a-mediated vasoconstriction and V2-mediated water retention with hyponatremia [12]. Counter-regulatory mechanisms oppose these effects but become blunted as disease progresses. Natriuretic peptides normally promote natriuresis, vasorelaxation, and inhibition of RAAS and fibrosis, but their renal actions become ineffective in advanced HF [13]. Endothelial nitric oxide is diminished by oxidative stress, and other vasodilatory mediators such as bradykinin, adrenomedullin, and apelin provide only limited protection. The imbalance between persistent neurohormonal activation and exhausted counter-regulation drives fluid retention, increased afterload, and progressive remodeling—processes that mechanical circulatory support aims to interrupt by restoring perfusion and reducing myocardial stress [15]. 2.3. HEMODYNAMIC EFFECTS OF MECHANICAL CIRCULATORY SUPPORT DEVICES: MCS devices aim to stabilize hemodynamics, unload the failing ventricle(s), maintain or augment cardiac output, and preserve end-organ perfusion [11,16]. These goals are achieved by assisting or entirely bypassing the native ventricle in propelling blood forward [17]. Depending on clinical needs and the specific ventricle involved, MCS devices can provide either partial or full support and can be used for either temporary or long-term purposes. Short-term MCS options include the intra-aortic balloon pump (IABP), left-sided percutaneous mechanical circulatory assist devices such as Impella (Abiomed, Danvers, MA, USA) and TandemHeart (LivaNova, London, UK), percutaneous right ventricular assist devices such as Protek Duo (LivaNova, London, UK) or Impella RP Flex (Abiomed, Danvers, MA, USA), and veno-arterial extracorporeal membrane oxygenation (VA-ECMO). Long-term MCS options primarily consist of left ventricular assist devices (LVADs) [18]. An overview of hemodynamic characteristics and clinical indications of different MCS devices is provided in Table 3. While the primary goal of MCS is to enhance end-organ perfusion and decrease myocardial oxygen demand, the resulting alterations in hemodynamic parameters compared with normal physiology are illustrated in Table 4 [19]. The following sections outline the mechanisms, physiological effects, and clinical implications of the respective devices. 2.4.1. INTRA-AORTIC BALLOON PUMP (IABP): Mechanism: A balloon catheter placed in the descending aorta inflates during diastole and deflates just before systole to augment diastolic blood pressure and provide systolic unloading [20]. Physiological Effects: IABP provides modest hemodynamic support primarily by increasing diastolic blood pressure, which significantly enhances coronary perfusion pressure. Additionally, the decreased afterload during systole reduces left ventricular end-diastolic pressure (LVEDP) and left ventricular wall stress, potentially enhancing cardiac output (CO) (0.5–1 L/min) and decreasing myocardial oxygen demand [21] as shown in Figure 1. However, clinical studies have shown mixed results. For instance, in patients with chronic heart failure, IABP improved mixed venous oxygen saturation, suggesting increased tissue perfusion [22]. Conversely, in patients with acute heart failure secondary to myocardial infarction, IABP did not significantly improve cardiac output, pulmonary capillary wedge pressure (PCWP), or mean arterial pressure (MAP) [23]. Thus, the hemodynamic benefits of IABP may vary depending on the underlying clinical scenario and patient characteristics. Clinical Management: The IABP is most effective in patients with partially preserved cardiac function, commonly serving as adjunctive rather than primary support. It has limited efficacy in severe cardiogenic shock when compared to other MCS options [24]. Although no MAP target has been definitively established, maintaining a MAP of approximately 65 mmHg has been suggested as a general guideline for patients with cardiogenic shock requiring temporary MCS [7]. 2.4.2. IMPELLA (LEFT-SIDED): Mechanism: A microaxial catheter-based pump positioned across the aortic valve directly unloads the left ventricle, pumping blood into the ascending aorta. Physiological Effects: The Impella device reduces LV end-diastolic pressure and volume, which decreases preload and pulmonary congestion because of LV unloading as shown in Figure 1. The Impella also lowers myocardial oxygen demand through reduced wall stress and increases cardiac output (up to approximately 3.5 L/min for Impella CP, 5.0 L/min for Impella 5.0, and 5.5 L/min for Impella 5.5). Clinical Management: A left-sided Impella is beneficial in severe LV dysfunction, high afterload conditions, or high-risk percutaneous coronary interventions (PCI). It improves coronary and systemic perfusion and may promote myocardial recovery but does not support RV function; thus, caution is needed in biventricular failure [6,25]. Hemodynamic management should target a MAP above 65 mmHg, with careful preload optimization to prevent suction events [26]. Preload optimization can be guided by echocardiography or invasive hemodynamic monitoring, including the calculation of LVEDP, which is available specifically in the Impella CP and 5.5 models with SmartAssist. Afterload must be managed judiciously, potentially using vasodilators if SVR is elevated, which may limit the effectiveness of the device. 2.4.3. TANDEMHEART: Mechanism: The TandemHeart device, placed transeptally via a transvenous approach, functions as a left atrial-to-femoral arterial bypass circuit, withdrawing oxygenated blood from the left atrium and returning it directly to systemic circulation using an ECMO pump. Physiological Effects: The TandemHeart device reduces LV preload, wall stress, and LVEDP, as well as PCWP. Unlike ECMO, it offers the advantage of significantly decreasing LV preload through more direct emptying, which is beneficial for patients with pulmonary congestion and edema. However, it does require intrinsic LV ejection and may increase LV afterload. Without adequate unloading strategies, there is a risk of LV distention. The pressure-volume loop typically narrows due to reduced LV end-diastolic volume and increased end-systolic volume [19] as shown in Figure 1. Clinical Management: The TandemHeart provides substantial support (up to 5 L/min) but requires invasive procedures, such as atrial septostomy, which limits its routine use [27,28]. 2.4.4. RIGHT VENTRICULAR ASSIST DEVICE (RVAD): Mechanism: RVADs draw blood from the right atrium or ventricle and pumps it into the pulmonary artery, thereby supporting failing right ventricular function (e.g., isolated RV failure, post-cardiotomy RV failure, RV failure post LVAD implantation, or as a bridge to transplantation). Surgical RVAD options include the CentriMag (Abbott Laboratories, Chicago, IL, USA), while percutaneous alternatives include the Protek Duo and the Impella RP. The CentriMag and Protek Duo configurations allow for oxygenator placement for patients with concomitant respiratory failure. Physiological Effects: RVADs increase pulmonary blood flow, improving LV preload and systemic output if LV function or LV support is adequate. They also reduce central venous pressure (CVP) and thus decrease hepatic and renal congestion. RV unloading decreases RV wall tension and myocardial oxygen consumption [29]. Clinical Management: Patients initially supported with an RVAD may require careful fluid resuscitation to optimize preload, transitioning to diuresis once stable. RV afterload must be optimized by using pulmonary vasodilators such as milrinone [30]. The effects of inhaled pulmonary vasodilators (e.g., epoprostenol or nitric oxide) should be carefully considered in the presence of significant LV dysfunction, as this may adversely impact LV performance. Close monitoring of RVAD flow (typically 3–5 L/min), echocardiographic assessment, and hemodynamic parameters (CVP, pulmonary artery pressures) are critical. Proactive management of potential complications, including hemolysis, thrombosis, bleeding, and infections, is important [29,30]. 2.4.5. VENO-ARTERIAL (VA-ECMO): Mechanism: VA-ECMO withdraws venous blood through a venous cannula, externally oxygenates it, and returns it to systemic circulation via arterial cannulation (commonly the femoral artery in peripheral VA-ECMO), effectively bypassing heart and lung functions [31,32]. Central cannulation, typically performed via the ascending aorta, is used in the surgical settings and provide direct access for arterial return. Axillary artery cannulation is a safe alternative, often used in patients with peripheral vascular disease or limited femoral access. Both central and axillary cannulation deliver antegrade (forward) flow, whereas femoral cannulation produces retrograde flow, with important physiological consequences. Peripheral cannulation via the femoral artery carries the riks of limb ischemia and may require placement of a distal perfusion catheter to ensure adequate limb flow. Close monitoring is important and should include physical examination, Doppler signals, or near-infrared spectroscopy. Physiological Effects: VA-ECMO significantly increases systemic perfusion and decreases RV preload but concurrently increases LV afterload due to retrograde arterial flow with peripheral cannulation as shown in Figure 1. The increased afterload elevates LVEDP, left atrial pressure, and PCWP, raising the risk of LV distention, higher myocardial oxygen demand, and pulmonary edema. This increased LV afterload, and the risk of LV distension are mitigated when arterial cannulation is performed centrally or axillary rather than femorally. Additionally, the LV preload is not necessarily reduced; the LV can still fill from several sources, including thebesian vessels, bronchial vessels, right heart blood not drained by the venous cannula, or retrograde flow from aortic regurgitation [33]. While inotropic support can provide “chemical unloading”, mechanical adjunctive measures such as IABP, left ventricular Impella, surgical ventricular venting, or atrial septostomy are often necessary to unload the LV [34]. However, data comparing unloading methods are limited; observational studies suggest early Impella unloading may improve outcomes [35], but no randomized trials have definitively established superiority [34]. Despite these potential risks, increased afterload from ECMO can improve coronary perfusion, increase oxygen supply, and correct metabolic derangements due to hypoperfusion, potentially aiding LV function. Clinical Management: VA-ECMO can be lifesaving in refractory cardiogenic shock or cardiac arrest (extracorporeal cardiopulmonary resuscitation—ECPR—scenarios) [36]. Effective management involves ensuring adequate venous drainage, promptly correcting hypovolemia, and optimizing LV ejection by utilizing inotropes or reducing afterload with vasodilators when clinically appropriate. Maintaining a pulse pressure above 10 mmHg or a CO approximately at least equal to one-fifth of ECMO flow is recommended to facilitate LV ejection, minimizing risks of LV distention and thrombus formation in the aortic root [37]. A significant complication associated with peripheral VA-ECMO is Harlequin syndrome (also known as north–south syndrome), characterized by differential oxygenation [38]. This occurs in the setting of concomitant respiratory failure when poorly oxygenated blood propelled by native LV function mixes with oxygenated retrograde ECMO blood distal to critical proximal aortic branches, leading to hypoxia of cerebral and coronary arteries. Monitoring with a right radial arterial line is recommended to promptly detect this complication. A narrow pulse pressure suggests that the mixing point is closer to the aortic root with lower risk, whereas a larger pulse pressure indicates potential for a more distal mixing point and higher risk. Further, arterial blood gas obtained from a right arterial line would indicate hypoxemia despite effective oxygenation by the ECMO circuit. Management strategies include optimizing ventilation, considering more central arterial cannulation, reducing LV ejection with negative inotropes, or employing hybrid cannulation strategies (e.g., veno-arterial-venous ECMO) [33]. Axillary or central arterial cannulation may improve upper-body oxygenation compared with femoral retrograde flow; however, these strategies alone do not prevent pulmonary edema, which is primarily related to LV distention and elevated afterload [31,32]. 2.4.6. DURABLE LEFT VENTRICULAR ASSIST DEVICE (LVAD): Mechanism: LVADs withdraw blood from the LV and pump it directly into the ascending aorta, thereby unloading the LV, reducing LVEDP, decreasing myocardial oxygen demand, and improving systemic perfusion. The most widely used durable LVAD currently available is the continuous-flow HeartMate 3 (Abbott Laboratories, Chicago, IL, USA) [39]. Physiological Effects: Adequate LVAD function depends on optimized preload (preload dependence). Hypovolemia and/or RV dysfunction may reduce LV and device filling and thus forward flow. At its worst, this may lead to suction events when the ventricle collapses around the device inflow and the pump then cannot flow. Elevated SVR increases the pump work required for forward flow, potentially diminishing LVAD performance (afterload sensitivity) [40]. Continuous-flow LVADs typically produce continuous blood flow without the need for LV ejection, leading to a triangular and leftward-shifted pressure-volume loop as shown in Figure 1 [19]. Increased flow and thus arterial pressure improve oxygenation, myocardial perfusion, and correct metabolic derangements. However, continuous flow often reduces or eliminates aortic valve opening, increasing the risk of blood stasis in the aortic root and potential thrombus formation. Continuous forward flow may also exacerbate aortic regurgitation, creating a feedback loop that decreases systemic output and elevates LV filling pressures. LV unloading reduces myocardial wall stress and oxygen consumption, potentially promoting myocardial recovery [19]. Typically, left atrial and pulmonary pressures decrease, improving pulmonary congestion and dyspnea. Nevertheless, the reduced pulsatility of the aortic pressure waveform may adversely impact baroreceptor reflexes and end-organ perfusion, particularly affecting kidney and brain function. Clinical Management: The target MAP for LVAD-supported patients is generally 70–80 mmHg [41]. Adequate preload is essential for optimal LVAD function, but attention must be given to avoiding RV failure precipitated by hypervolemia or increased RV afterload. CVP monitoring, ideally maintained below 15 mmHg, is helpful to guide fluid management, diuresis, and inotropic support if needed [42]. Echocardiography plays a key role in assessing LV unloading, aortic valve function, inflow and outflow position, and guiding LVAD speed and flow adjustments, especially in patients with elevated CVP due to positive pressure ventilation. Afterload should be carefully managed, using vasopressors for reduced afterload and cautious use of vasodilators when afterload is elevated. Maintenance of aortic valve opening is important for reducing thrombotic risks and improving outcomes [43]. Should a suction event occur, management involves reducing pump speed initially, followed by targeted treatment of the underlying cause (e.g., hypovolemia, RV failure, cardiac tamponade) [44]. Individualized anticoagulation plans are necessary based on patient-specific bleeding and thrombosis risks [45]. Clinicians must remain vigilant for complications including gastrointestinal bleeding (associated with acquired von Willebrand deficiency and arteriovenous malformations due to laminar flow), thrombosis, hemolysis, and infections [40]. 2.4.7. BIVENTRICULAR ASSIST DEVICE (BIVAD): Mechanism: BiVADs typically consist of combined LVAD and RVAD systems and employed to simultaneously support both the left and right ventricles in patients with severe biventricular failure. Durable BiVAD systems commonly involve two separate surgically implanted centrifugal flow pumps, such as the HeartMate 3, each dedicated to supporting one ventricle. Temporary BiVAD options include a serial surgical RVAD and left-sided Impella device or percutaneous configuration of biventricular catheter-based pumps, like the Impella or ProtekDuo, with the latter providing more rapid deployment and stabilization [46,47]. Physiological Effects: The optimal function of BiVADs requires meticulous coordination of device outputs to ensure balanced ventricular unloading and adequate systemic and pulmonary perfusion. Proper matching of RVAD and LVAD flows is essential; if RVAD flow exceeds LVAD flow, it may precipitate pulmonary overcirculation, which results in pulmonary edema due to increased pulmonary blood volume and elevated pulmonary pressures. Conversely, excessive LVAD flow relative to RVAD output may lead to systemic hypoperfusion and inadequate preload to the left ventricle, compromising end-organ perfusion. Additionally, imbalanced support between the two ventricles can result in interventricular septal shifts, distorting ventricular geometry and adversely impacting ventricular contractility and overall cardiac efficiency. Clinical Management: Effective BiVAD management thus requires careful coordination of preload optimization, afterload modulation, and device flow coordination. Management of BiVAD-supported patients requires continuous and vigilant monitoring. This involves both invasive hemodynamic assessments (such as pulmonary artery catheters) and non-invasive modalities (echocardiography). Frequent adjustments to device flow settings based on these comprehensive evaluations in a dynamic patient environment are necessary to maintain optimal hemodynamics. Regular echocardiographic assessments help detect complications such as septal malposition, ventricular distension, or cannula malposition, and guide flow adjustments to restore balanced support [48]. Anticoagulation strategies for the major MCS devices are outlined in Table 5 [49,50]. 2.5. VENTRICULAR-ARTERIAL COUPLING, OXYGEN DELIVERY, AND METABOLIC-NEUROHORMONAL EFFECTS DURING MCS: Ventricular-arterial coupling refers to the relationship between ventricular contractility and arterial load, a critical determinant of cardiovascular efficiency. MCS significantly modifies this relationship. For example, LVADs typically improve LV-arterial coupling by reducing afterload and enhancing systemic perfusion, thus decreasing myocardial workload and improving ventricular efficiency. Conversely, with RVADs, RV-arterial coupling heavily depends on pulmonary vascular resistance (PVR). Elevated PVR can limit RVAD effectiveness, necessitating adjunct therapies such as pulmonary vasodilators (e.g., inhaled nitric oxide) to optimize RV support and overall hemodynamic stability [19]. MCS devices enhance tissue oxygen delivery (DO2) primarily through improved systemic perfusion. In shock states, impaired DO2 leads to anaerobic metabolism and subsequent lactic acidosis. Restoration of adequate DO2 by MCS facilitates aerobic metabolism and lactate clearance. Continuous monitoring of markers such as lactate, central venous oxygen saturation (ScvO2), and mixed venous oxygen saturation (SvO2) is crucial for guiding clinical management and assessing therapeutic effectiveness [51]. Persistent neurohormonal activation (e.g., RAAS, sympathetic nervous system) and microcirculatory dysfunction may continue despite the normalization of macrocirculatory parameters (e.g., MAP and CO) with MCS. MCS may improve end-organ function by increasing perfusion. However, continuous-flow LVADs can induce non-physiologic flow patterns, potentially affecting endothelial function, impairing baroreceptor reflexes, aut
PMID:41096015 | 1 | 41096015 | PMC12524590 | 10.3390/jcm14196935 | fetched |  | oregulation and renal perfusion [51]. Applying these physiological principles to clinical decision-making is important. Figure 2 provides a practical algorithm to guide the selection of temporary MCS devices based on the severity and type of ventricular dysfunction. An overview of long-term MCS selection pathways is provided in Figure 3. 2.6. CLINICAL ALGORITHMS FOR MCS SELECTION: The application of MCS in clinical practice requires structured decision pathways. The choice of temporary MCS in cardiogenic shock needs to consider factors such as ventricular involvement, prognosis, anticipated duration of support, and feasibility, as shown in Figure 2 [7,18]. Pathways for durable MCS in advanced heart failure must take into account eligibility for transplantation, contraindications to long-term support (e.g., severe comorbidities, irreversible end-organ dysfunction, or active infection), and therapeutic options such as bridge-to-transplant, destination therapy, or palliative care, as shown in Figure 3 [52,53]. 3.1. IABP: First introduced in 1968 [54], the IABP has been widely used for cardiogenic shock and perioperative support [55,56]. It is safe, cost-effective, and percutaneously inserted, with major complications occurring in <1% of cases [56,57]. Although initially recommended as Class I therapy, the IABP-SHOCK II trial demonstrated no survival benefit in myocardial infarction–related cardiogenic shock [58,59], leading to guideline downgrading to Class III (no benefit) [60,61]. A meta-analysis suggested potential benefit when used preoperatively in high-risk CABG patients [62], but overall clinical practice has shifted toward newer MCS devices. 3.2. IMPELLA: The Impella, derived from the Hemopump (1985), received CE approval in 2005 and FDA approval in 2008 [63]. It provides hemodynamic support in high-risk PCI, cardiogenic shock, LV failure, RV failure (Impella RP Flex), and perioperative settings (e.g., during ventricular tachycardia ablation, high-risk PCI, post-cardiopulmonary bypass, or post-transcatheter aortic valve replacement). It can also serve as a bridge to transplant or for LV venting during VA-ECMO [64]. Smaller devices (e.g., Impella CP) are placed percutaneously, while Impella 5.5 requires surgical transaxillary implantation, allowing longer-term support and patient mobility. Complications include vascular injury (7–8%), bleeding (15%), hemolysis (~10%), and, with the transaxillary approach, up to 17% serious complications and ~50% in-hospital mortality in high-risk patients [65,66]. Evidence for outcomes is mixed. The Impella 2.5 provides limited hemodynamic support and has not shown a survival benefit in cardiogenic shock [66]. In contrast, larger devices (5.0/5.5) demonstrate improved outcomes, with ~70% one-year survival in post-cardiac surgery shock [67,68]. The DanGer Shock trial reported a 12.7% absolute reduction in 180-day mortality with Impella CP in STEMI-related shock, although this was associated with higher complication rates, including bleeding, limb ischemia, and the need for renal replacement therapy [6]. Earlier studies such as PROTECT II found no significant advantage of Impella 2.5 over IABP in high-risk PCI [69], and subsequent analyses also showed no 30-day mortality difference between Impella 2.5 and IABP in myocardial infarction–related shock, despite higher complication rates with Impella [70]. To date, no direct comparative trials exist between IABP and larger Impella devices. 3.3. VA-ECMO: VA-ECMO rapidly stabilizes patients with cardiac arrest, refractory cardiogenic shock, post-cardiotomy failure, or acute intra-procedural collapse, serving mainly as a temporary bridge to recovery or definitive therapy [71]. Complications of prolonged use include bleeding, systemic inflammation, limb ischemia, thrombotic events, and LV distension, often requiring adjunctive unloading with IABP or Impella [72,73]. Evidence for ischemia-related cardiogenic shock is mixed. The EURO SHOCK trial (terminated early due to COVID-19) suggested possible survival benefit, though with increased bleeding and vascular events [74]. In contrast, the larger ECLS-SHOCK trial showed no improvement in 30-day or 12-month mortality and confirmed higher complication rates [75]. A recent meta-analysis likewise found no overall survival benefit from early percutaneous MCS in AMI-related shock, except in a subgroup of STEMI patients without hypoxic brain injury [9,72,73]. Registry and observational data highlight potential benefits in specific populations: survival of 34–100% in pulmonary embolism [71], 61% in myocarditis [76], and procedural stability during ventricular tachycardia ablation [77]. Meta-analyses comparing Impella and VA-ECMO show similar mortality but fewer bleeding and stroke complications with Impella [78]. Combined use of Impella and VA-ECMO may improve short-term outcomes but is associated with higher risks of hemolysis, ischemia, and renal failure requiring dialysis [79]. 3.4. DURABLE LVAD: Durable LVADs are indicated for advanced heart failure (NYHA III–IV) refractory to medical therapy, serving as a bridge to transplant, bridge to candidacy, or destination therapy [80,81]. The MOMENTUM 3 trial reported HeartMate 3 survival rates of 79% at 2 years and 58% at 5 years [82], consistent with ELEVATE registry findings (74.5% and 63.3%) [83]. Common complications include bleeding, driveline infections, and progressive aortic regurgitation (≈30%) [84]. Right heart failure occurs in 20–40% of patients, requiring RV support in ~5% [85]. While pump thrombosis is less frequent with HeartMate 3, gastrointestinal bleeding, arrhythmias, and stroke remain important concerns [84]. 3.5. BIVAD: Patients with chronic biventricular failure remain difficult to manage, and evidence is limited. Off-label use of dual HeartMate 3 pumps has emerged as a strategy for biventricular support. In a multicenter series (14 patients), mean support duration was 266 days, with most discharged home [46]. A larger cohort (38 patients) reported one-third mortality during support, while nine patients were bridged to transplant and 18 maintained on support for ~5 months [86]. Complications included device malfunction (25%), bleeding, neurologic, and renal dysfunction. A concise summary of the key evidence and clinical indications for the MCS devices discussed is presented in Table 6. 4.1. PULMONARY ARTERY CATHETER (PAC): PAC monitoring can provide valuable hemodynamic information in patients with cardiogenic shock requiring MCS by helping to categorize the type and severity of shock, guide escalation or de-escalation of support, and assess adequacy of resuscitation [87,88]. Randomized data establishing a treatment algorithm based on PAC findings in cardiogenic shock are lacking, and much of the current evidence is observational. In one observational study, PAC insertion prior to MCS initiation was associated with reduced mortality; however, the strongest predictors of outcome were mean arterial pressure and central venous pressure, which are not specific to PAC [89]. More comprehensive hemodynamic profiling (including right atrial pressure, pulmonary artery pressures, pulmonary capillary wedge pressure, cardiac output, and PA oxygen saturation) showed potential prognostic value, but the overall role of PAC in guiding therapy remains uncertain and should be interpreted with caution [89]. In patients already on VA-ECMO, increased PA pulse pressure compared to pre-ECMO values has been correlated with improved survival, likely reflecting preserved RV function [90]. Ongoing prospective trials such as PACCS (ClinicalTrials.gov Identifier: NCT05485376), which is testing whether early PAC-guided management in cardiogenic shock due to acutely decompensated heart failure reduces in-hospital mortality, may further clarify the role of PAC monitoring. Importantly, this trial excludes patients receiving short-term mechanical support and does not prespecify treatment protocols. 4.2. CONTINUOUS CARDIAC OUTPUT AND SVO2: Continuous monitoring of CO and SvO2 using a PAC provides critical information on systemic perfusion and tissue oxygenation during MCS. Accurate assessment of SvO2 helps evaluate the balance between DO2 and O2 consumption, guiding adjustments in device settings and therapeutic management. Specifically, in patients supported with VA ECMO, SvO2 measurements accurately reflect systemic perfusion status and are routinely used to monitor the adequacy of circulatory support [91]. 4.3. ECHOCARDIOGRAPHY: Echocardiography, including transthoracic (TTE) and transesophageal (TEE) approaches, is critical in managing patients with MCS, providing real-time assessment of cardiac anatomy, ventricular function, device positioning, and hemodynamic status. TTE is typically the initial modality due to its non-invasive nature and convenience for rapid bedside assessment; however, its image quality can be limited by factors such as body habitus and acoustic shadowing from implanted devices. TEE is frequently used in the perioperative and intensive care settings for accurate device placement, real-time monitoring, and early detection of complications, especially when TTE imaging quality is suboptimal [92]. Further, echocardiography has a specific utility for individual MCS devices [93]. For Impella devices, it is used to confirm appropriate positioning across the aortic valve, verify effective left ventricular unloading, and identify complications such as valve damage, ventricular wall injury, or device migration. For TandemHeart and ProtekDuo devices, echocardiography, particularly TEE, is used to guide critical steps such as transseptal puncture (TandemHeart), optimal cannula positioning, and effective ventricular decompression. It is also used to confirm correct inflow cannula placement in the left atrium (TandemHeart) or the pulmonary artery (ProtekDuo). In patients receiving VA-ECMO, particularly in the setting of ECPR, echocardiography is valuable for excluding contraindications such as aortic dissection and severe aortic insufficiency, before implantation. Post-initiation, it is important for detecting complications, including left ventricular distension, inadequate ventricular ejection, cannula malposition, and thrombus formation; however, imaging can be challenging due to acoustic shadowing from ECMO cannulas [94]. During weaning from VA-ECMO, echocardiographic parameters predictive of successful decannulation include increased aortic valve opening, improved left ventricular ejection fraction (LVEF ≥20–25%), left ventricular outflow tract velocity-time integral (LVOT VTI >10 cm), adequate RV function (RV ejection fraction ≥25%, and Tricuspid Annular Plane Systolic Excursion [TAPSE] ≥19 mm) [95]. Therefore, echocardiography is indispensable for safe and effective placement, management, and weaning of patients receiving MCS. 4.4. ARTERIAL WAVEFORM ANALYSIS: Arterial waveform analysis provides valuable hemodynamic information during MCS. In the context of IABP management, waveform analysis helps evaluate ventricular unloading and device performance. In patients receiving Impella support, internal waveform data, including motor-current and pressure signals from the device controller, assist with correct positioning and early detection of complications such as suction events, although standard arterial pressure monitoring remains necessary. During VA-ECMO support, arterial waveform analysis is especially beneficial for monitoring for recovery and during weaning phases as the reappearance of pulsatile waveforms typically indicates improved myocardial function and sufficient native CO [19]. Commercially available monitoring systems analyze arterial waveforms from standard arterial lines and derive parameters such as SV, CO, stroke volume variation (SVV), and pulse pressure variation (PPV). Despite their established utility in the perioperative and critical care environments, the accuracy and applicability of these devices during MCS are still unknown but likely limited by reduced arterial pulsatility and variable SVR [96]. 4.5. END-TIDAL CARBON DIOXIDE (ETCO2): End-Tidal Carbon Dioxide (EtCO2) monitoring provides a non-invasive assessment of pulmonary blood flow and indirectly reflects native CO, making it particularly valuable during ECMO support and weaning phases. An increase in EtCO2 suggests improved native cardiac function, as increased pulmonary blood flow results in greater CO2 clearance. A study involving 37 VA ECMO patients demonstrated that an EtCO2 rise of at least 5 mmHg over 12 h reliably predicted successful ECMO weaning [97]. Thus, EtCO2 monitoring serves as a valuable adjunct to invasive hemodynamic assessment, facilitating clinical decision-making for optimal timing of ECMO weaning. 4.6.1. VA-ECMO PARAMETERS: FLOW RATES AND OXYGENATION INDICES: ECMO flow serves as a surrogate for CO, with recommended target flows between 50 and 75 mL/kg/min for adult patients. Larger patients may require additional venous drainage cannulas to achieve adequate perfusion. Gas exchange during ECMO relies on a gradient created by fresh gas, termed “sweep gas,” which removes CO2 and delivers O2. Oxygenation improves with increased blood flow or higher fractions of delivered oxygen, independent of sweep gas flow, because more blood is exposed to the oxygenator membrane. Conversely, CO2 elimination directly correlates with sweep gas flow; increasing sweep gas enhances CO2 removal, analogous to increasing minute ventilation in patients without ECMO support [98]. 4.6.2. LVAD PARAMETERS (HEARTMATE 3): PUMP SPEED, FLOW, AND PULSATILITY INDEX: HeartMate 3 is the dominant device for durable LVAD support. Its pump speed typically ranges between 3000 and 9000 revolutions per minute (RPM), but it is commonly maintained above 4000 RPM. Though a continuous flow device, the HeartMate 3 employs intrinsic pulsatility through periodic decreases and increases in device speed every two seconds to improve flow dynamics and decrease the risk of thrombosis. LVAD flow depends on both RPM and afterload; flow decreases as afterload increases. By contrast, reduced preload may trigger a “low-speed limit,” known as a “suction event.” Flow estimates are derived from pump power consumption; due to the linear relationship between flow and power at a fixed speed, factors like pump thrombosis that increase power can produce falsely elevated flow readings. The pulsatility index (PI) reflects increased pump flow during ventricular systole and typically remains below 10 in optimal conditions. Lower PI values indicate reduced ventricular activity, while higher values suggest increased pulsatility and thus improved cardiac function. A stable PI at rest is expected, and an abrupt PI drop can signal reduced preload [40]. 5.1. PHYSIOLOGICAL ADAPTATION DURING WEANING FROM MCS: One of the key physiologic changes in MCS for any indication is ventricular offloading to decrease wall stress and myocardial oxygen demand [99]. This is theoretically more notable in central cannulation and temporary percutaneous VADs, where unloading is much more complete due to more proximal drainage [100]. In contrast, patients on peripheral VA ECMO may face increased LV afterload, particularly at higher flows, due to the retrograde nature of perfusion [100]. As such, the weaning from MCS innately requires reloading of a recently stressed ventricle, and thus liberation from MCS hinges on tolerance of this process, particularly for the right ventricle [100]. There are currently no standardized protocols regarding the timing or strategies for tMCS weaning. Assuming the underlying cause has been treated, improved, or resolved, weaning should ideally be initiated within 48–72 h of MCS initiation with daily reassessment. Figure 4 illustrates a structured algorithm for temporary MCS weaning. Some patients status post LVAD placement have functional myocardial recovery and reverse structure remodeling due to the long-term offloading of the left ventricle allowing the previous “stretch” to be undone [101]. The RESTAGE-HF trial demonstrated that nearly half of patients achieved improvement such that LVAD explantation was accomplished and recovered function was sustained for a year [102]. This physiologic change is the premise behind BTR pathways previously discussed. The molecular pathways to explain this are being investigated [101]. Recently, this finding was explored in temporary MCS modalities including prolonged use of an Impella device via the PROPELLA trial which identified disease-modifying effects of Impella which predisposed myocardial recovery in patients with fulminant myocarditis [103]. 5.2. ETHICAL CONSIDERATION IN MCS: MCS is often initiated rapidly in acutely decompensating patients, forcing clinicians, patients, and families to make life-or-death decisions under conditions of uncertainty and incomplete information. When caring for patients receiving MCS, clinicians should apply the principles of clinical bioethics and carefully balance the potential benefits of support against the risks of futility, prolonged suffering, and poor quality of life. 6. EMERGING ROLE OF ARTIFICIAL INTELLIGENCE IN MCS: Artificial Intelligence (AI) refers to the ability of computer systems to learn from data and perform tasks that typically require human intelligence, such as decision-making and problem-solving. Machine Learning (ML), a subset of AI, involves training these systems to recognize complex patterns and accurately predict outcomes directly from data. Recently, AI prediction models have significantly advanced cardiac care, with applications in echocardiography interpretation, electrocardiogram (EKG) analysis, and heart transplant outcome predictions [114]. In the context of MCS, AI can utilize extensive clinical datasets to identify optimal candidates, anticipate potential complications such as right ventricular failure, and determine the most appropriate type of mechanical support. Post-implantation, AI-driven algorithms can dynamically adjust MCS parameters (e.g., LVAD speed, flow rates, pulsatility index) based on real-time patient data, potentially optimizing clinical management [115,116]. These concepts remain largely proof-of-concept, demonstrating feasibility but not yet established in clinical practice. Some applications are beginning to approach clinical relevance. Recent studies using Bayesian and decision tree models derived from the INTERMACS registry have predicted short- and long-term outcomes following LVAD implantation [117]. Similarly, advanced ML techniques, such as convolutional neural networks and decision tree model, have been applied to echocardiographic analysis for assessment of RV function, improving patient selection and perioperative planning [118,119]. Additionally, ML models analyzing electronic health records and wearable devices have shown potential for early detection of clinical deterioration, facilitating timely interventions, and reducing hospital readmissions [120]. These applications represent important translational steps, though widespread clinical adoption will require prospective validation. Despite these promising advances, significant challenges remain, including issues related to data integration, model interpretability, and resource availability. At present, the majority of AI work in MCS should be considered exploratory. Continued research and rigorous clinical trials are needed to refine these technologies, validate their effectiveness, and ultimately integrate them safely into routine patient care. 7. CONCLUSIONS: MCS is a rapidly evolving field requiring a thorough understanding of its cardiovascular interactions and careful selection of monitoring strategies. Given the complexity of MCS devices, individualized patient management, accurate interpretation of monitoring data, and attention to ethical considerations are essential. Future research should focus on addressing device-specific limitations, clarifying their clinical utility, refining monitoring methods, and ensuring safe integration into practice. Emerging technologies like AI require further validation before routine clinical use.
PMID:40634051 | 0 | 40634051 |  | 10.1016/j.tvir.2025.101042 | fetched | {"title": "ECMO, VAD, Impella, and IABP: Primer on Mechanical Circulatory Support Devices.", "authors": ["Kruse J", "Ranade M"], "journal_abbrev": "Tech Vasc Interv Radiol", "year": 2025, "volume": "28", "issue": "2", "pages": "101042", "doi": "10.1016/j.tvir.2025.101042", "pmid": "40634051", "pmcid": null, "pub_types": ["Journal Article", "Review"]} | ECMO, VAD, Impella, and IABP: Primer on Mechanical Circulatory Support Devices. Vascular and interventional physicians are routinely caring for some of the sickest patients in the hospital, including those with cardiopulmonary collapse. A foundational comprehension of mechanical circulatory devices is necessary in these situations. Within this article, we will review the indications, procedural steps, management, clinical follow up, and outcomes of 4 types of mechanical circulatory support devices: Extracorporeal membrane oxygenation (ECMO), ventricular assist devices (VAD), Impella, and intra-aortic balloon pumps (IABP). ECMO is a large-bore cannulation system serving patients in respiratory and/or combination of cardiorespiratory collapse. It requires constant monitoring and adjustment with a wide range of benefits in appropriately selected patients in the hospital. VADs are surgically-implanted devices for more long-term ventricular support, often extending the survival of end-stage heart failure or bridging to transplant. Impella and IABPs, by comparison, are smaller intravascular devices that aid in short-term ventricular off-loading and improve perfusion. They provide hemodynamic stabilization in the setting of acute cardiac insult or surgery, prior to either organ recovery or transition to more definitive therapy. All of these devices come with unique considerations and potential complications. This article will serve as a guide in evaluating patients for device candidacy, procedural placement, troubleshooting, management in the perioperative setting, complication surveillance, and clinical outcomes/follow-up.
PMID:8573871 | 0 | 8573871 |  | 10.1097/00002480-199507000-00077 | fetched | {"title": "Effects of timing on ventriculoarterial coupling and mechanical efficiency during intraaortic balloon pumping.", "authors": ["Sakamoto T", "Suzuki A", "Kazama S", "Komatsu S", "Sasaki S", "Shoji Y"], "journal_abbrev": "ASAIO J", "year": 1995, "volume": "41", "issue": "3", "pages": "M580-3", "doi": "10.1097/00002480-199507000-00077", "pmid": "8573871", "pmcid": null, "pub_types": ["Comparative Study", "Journal Article"]} | Effects of timing on ventriculoarterial coupling and mechanical efficiency during intraaortic balloon pumping. To evaluate the effects of deflation timing of an intraaortic balloon (IAB) on ventriculoarterial (VA) coupling and left ventricular (LV) mechanics, we assessed contractility: Emax (mmHg/ml); VA coupling: Ea/Ees; myocardial oxygen consumption: LV pressure-volume area (PVA); and energy efficiency: external work (EW)/PVA during IAB pumping, using four different timing intervals of IAB deflation in eight dogs weighing 19.8 +/- 3.8 kg. In the conventional timing group (G-conv), the IAB was deflated before LV systole. In the real time group (G-1, 2, 3), the IAB was deflated during isovolumic contraction (ICP), and the group was subdivided into three groups based on time of deflation: G-1 at the onset of ICP, G-2 at the midpoint of ICP, and G-3 at the end of ICP. Percent changes in values with IAB compared to those without assist were calculated. Based on the results, IAB deflation during ICP reduces LV end-systolic pressure and significantly optimizes VA coupling and LV efficiency during IAB support, compared with conventional timing.
PMID:23263334 | 0 | 23263334 |  | 10.1097/MAT.0b013e3182768ba9 | fetched | {"title": "Modeling the interaction between the intra-aortic balloon pump and the cardiovascular system: the effect of timing.", "authors": ["Schampaert S", "Rutten MC", "van T Veer M", "van Nunen LX", "Tonino PA", "Pijls NH", "van de Vosse FN"], "journal_abbrev": "ASAIO J", "year": 2013, "volume": "59", "issue": "1", "pages": "30-6", "doi": "10.1097/MAT.0b013e3182768ba9", "pmid": "23263334", "pmcid": null, "pub_types": ["Journal Article", "Research Support, Non-U.S. Gov't", "Validation Study"]} | Modeling the interaction between the intra-aortic balloon pump and the cardiovascular system: the effect of timing. Because of the large number of interaction factors involved, the effects of the intra-aortic balloon pump (IABP) have not been investigated deeply. To enhance its clinical efficiency and to better define indications for use, advanced models are required to test the interaction between the IABP and the cardiovascular system. A patient with mild blood pressure depression and a lowered cardiac output is modeled in a lumped parameter computational model, developed with physiologically representative elements for relevant components of circulation and device. IABP support is applied, and the moments of balloon inflation and deflation are varied around their conventional timing modes. For validation purposes, timing is adapted within acceptable ranges in ten patients undergoing IABP therapy for typical clinical indications. In both model and patients, the IABP induces a diastolic blood pressure augmentation as well as a systolic reduction in afterload. The support capabilities of the IABP benefit the most when the balloon is deflated simultaneously with ventricular contraction, whereas inflation before onset of diastole unconditionally interferes with ejection. The physiologic response makes the model an excellent tool for testing the interaction between the IABP and the cardiovascular system, and how alterations of specific IABP parameters (i.e., timing) affect this coupling.
PMID:36682892 | 0 | 36682892 |  | 10.1016/j.healun.2022.12.007 | fetched | {"title": "Effects of Intra-Aortic Balloon Pump delayed deflation timing on carotid blood flow and cardiac mechanics: An ultrasound and haemodynamic study.", "authors": ["Boccellino A", "Sacchi S", "Pazzanese V", "Calvo F", "Gramegna M", "Ingallina G", "Cappelletti A", "Baldetti L"], "journal_abbrev": "J Heart Lung Transplant", "year": 2023, "volume": "42", "issue": "4", "pages": "451-455", "doi": "10.1016/j.healun.2022.12.007", "pmid": "36682892", "pmcid": null, "pub_types": ["Journal Article"]} | Effects of Intra-Aortic Balloon Pump delayed deflation timing on carotid blood flow and cardiac mechanics: An ultrasound and haemodynamic study. Intra-Aortic Balloon Pump (IABP) efficacy is critically affected by the inflation/deflation timing. Balloon deflation may cause a sucking effect, and a steal phenomenon on carotid flow. Delaying IABP deflation reduces the degree of this flow reversal, but at the same time exposes patients to the risk of increased proto-systolic afterload with detrimental effects on the LV. The purpose of this study was to investigate the effects of a delayed IABP deflation timing on cerebral blood flow and LV hemodynamics, by means of simultaneous carotid artery ultrasonography, trans-thoracic echocardiography and central aortic pressure analysis. Delaying IABP deflation trigger to the beginning of QRS effectively increased the cerebral blood flow by 20%, mostly by reducing the reverse component flow caused by the diastolic balloon deflation. Extending the deflation to the early systole was safe and favourably impacted on cardiac mechanics, increasing CO by 15% without prolonging LV isovolumetric contraction and ejection phases.
PMID:27582283 | 0 | 27582283 |  | 10.1161/CIRCHEARTFAILURE.116.003474 | fetched | {"title": "Deflation Timing Influences Intra-Aortic Balloon Pump-Mediated Carotid Blood Flow Reversal: A Case Report.", "authors": ["Schutt RC", "Bhimiraj A", "Estep JD", "Guha A", "Trachtenberg BH", "Garami Z"], "journal_abbrev": "Circ Heart Fail", "year": 2016, "volume": "9", "issue": "9", "pages": null, "doi": "10.1161/CIRCHEARTFAILURE.116.003474", "pmid": "27582283", "pmcid": null, "pub_types": ["Case Reports", "Journal Article"]} | Deflation Timing Influences Intra-Aortic Balloon Pump-Mediated Carotid Blood Flow Reversal: A Case Report.
PMID:40495936 | 0 | 40495936 | PMC12150027 | 10.1093/ehjimp/qyaf067 | fetched | {"title": "Echocardiography for short-term mechanical circulatory support: a trans-Atlantic practical guide.", "authors": ["Soliman-Aboumarie H", "Vandenbriele C", "Baldetti L", "Balthazar T", "Ott S", "Hernandez-Montfort J", "Nicoara A"], "journal_abbrev": "Eur Heart J Imaging Methods Pract", "year": 2025, "volume": "3", "issue": "1", "pages": "qyaf067", "doi": "10.1093/ehjimp/qyaf067", "pmid": "40495936", "pmcid": "PMC12150027", "pub_types": ["Journal Article", "Review"]} | Echocardiography for short-term mechanical circulatory support: a trans-Atlantic practical guide. Short-term mechanical circulatory support (stMCS) devices are increasingly utilized for haemodynamic stabilization in patients with cardiogenic shock. Echocardiography plays a pivotal role across the continuum of stMCS use-from patient selection and device implantation to monitoring, troubleshooting, and weaning. This review provides a comprehensive, practical guide for clinicians on the echocardiographic assessment of commonly used stMCS devices, including intra-aortic balloon pump, Impella, ProtekDuo, and veno-arterial extracorporeal membrane oxygenation. We outline device-specific contraindications, key imaging views for guiding placement, and parameters for monitoring device performance and detecting complications. The paper also introduces structured echocardiographic criteria to support decision-making during weaning and explantation. Finally, we explore emerging tools such as speckle-tracking echocardiography, 3D imaging, and artificial intelligence that may further optimize stMCS management. Through an international, multidisciplinary collaboration, this guide aims to standardize echocardiographic practice in stMCS and improve clinical outcomes in critically ill cardiac patients. ABSTRACT: AbstractShort-term mechanical circulatory support (stMCS) devices are increasingly utilized for haemodynamic stabilization in patients with cardiogenic shock. Echocardiography plays a pivotal role across the continuum of stMCS use—from patient selection and device implantation to monitoring, troubleshooting, and weaning. This review provides a comprehensive, practical guide for clinicians on the echocardiographic assessment of commonly used stMCS devices, including intra-aortic balloon pump, Impella, ProtekDuo, and veno-arterial extracorporeal membrane oxygenation. We outline device-specific contraindications, key imaging views for guiding placement, and parameters for monitoring device performance and detecting complications. The paper also introduces structured echocardiographic criteria to support decision-making during weaning and explantation. Finally, we explore emerging tools such as speckle-tracking echocardiography, 3D imaging, and artificial intelligence that may further optimize stMCS management. Through an international, multidisciplinary collaboration, this guide aims to standardize echocardiographic practice in stMCS and improve clinical outcomes in critically ill cardiac patients. INTRODUCTION: Short-term mechanical circulatory support (stMCS) devices are increasingly utilized in acute cardiac care for the management of patients in cardiogenic shock (CS)1 by stabilizing haemodynamics and allowing for either myocardial recovery or long-term intervention decisions such as bridging to long-term (durable) ventricular assist devices or heart transplantation. Echocardiography is essential for the effective implementation and management of stMCS as it provides guidance during implantation, as well as information regarding device positioning, complications, and real-time longitudinal cardiac function assessment. However, despite its critical role, the use of echocardiography in stMCS management remains highly variable across institutions, with differing protocols, operator expertise, and access to advanced imaging techniques. This variability highlights an important unmet need for standardized approaches and training in echo-guided stMCS care. This review provides a practical overview of the role of echocardiography in stMCS in the initiation, guidance, troubleshooting, and weaning, while identifying best practices and highlighting areas for future research. OVERVIEW OF STMCS: Selecting the appropriate stMCS device requires an assessment of four key factors2: (i) severity of CS, (ii) univentricular vs. biventricular failure, (iii) presence of respiratory failure requiring oxygenation support, and (iv) anatomical or functional limitations or contraindications to certain devices. Echocardiography is a cornerstone in the evaluation of these patients and ensuring that device selection aligns with the underlying anatomical and haemodynamic profile. In the following sections, we will briefly review the use of echocardiography in the management of some of the most used stMCS devices3 (Figure 1). The table in the Supplementary data includes a summary of indications and specifics for each device. HAEMODYNAMIC PHENOTYPING: Echocardiography prior to stMCS is indispensable: it confirms cardiac dysfunction as the underlying driver of the circulatory shock; it allows morphologic and haemodynamic assessment of cardiac chambers—helping in the selection of the most appropriate stMCS device and identifies specific contraindications to device placement. In addition, a pre-implantation assessment provides a benchmark for comparison on repeated imaging during the clinical course of the patient, in order to track the subsequent haemodynamic trajectory. The role of non-invasive ‘echodynamic’ evaluation has shown to be reliable, except for the estimation of left atrial pressure, which remains difficult in the critically ill patients, especially in case of stMCS.4 Pre-implantation echocardiographic assessment should include a thorough exploration of the left and right ventricles (LV and RV) and atria, atrioventricular and semilunar valves, pericardial space, inferior vena cava (IVC), and lung fields. Notwithstanding LV ejection fraction (LVEF) as a straightforward measure of LV contractile function, we recommend to always assess LV outflow tract velocity-time integral (LVOT VTI) to non-invasively calculate LV stroke volume, as this represents a more sensitive and reliable gauge of LV ejection on repeated assessments and particularly once the stMCS has been deployed. Nevertheless, LVOT diameter should also be taken into consideration as it influences stroke volume (i.e. large LVOT diameter could compensate for reduced LVOT VTI achieving normal stroke volume). Moreover, whenever a stMCS is considered, assessment of the LV end-diastolic volume (LVEDV) and diameter (LVEDD) and of the LV diastolic function through trans-mitral pulsed wave Doppler and mitral annular tissue Doppler imaging is warranted. Right-sided chambers should also be explored to estimate longitudinal and radial RV function, RV outflow tract VTI, pulmonary artery (PA) pressure, and central venous pressure (CVP). Notably, assessment of right heart preload should be multiparametric, as positive end-expiratory pressure, pre-existing cor pulmonale, and severe tricuspid regurgitation (TR) can yield a dilated and non-collapsible IVC and a raised CVP despite being preload responsive. The recently proposed Venous Excess Ultrasound (VeXUS) score provides a multiparametric assessment of CVP and is more accurate than IVC dimension and collapsibility in identifying elevated CVP5–7 and predicts mortality in the acute heart failure setting.8 Lung ultrasound also plays a crucial role in assessing acute heart failure by detecting pulmonary congestion through B-lines, pleural effusions, and dynamic changes in lung aeration, providing valuable real-time information for patient management.9 Finally, given the strong interplay of LV and RV, the interventricular septum position and motion should always be evaluated. The degree of detail of any echocardiographic evaluation should be linked to the clinical stability of the patient, as timely and rapid stMCS may be required in highly unstable patients (e.g. Society of Cardiovascular Angiography and Interventions Class D to E patients). SUITABILITY FOR STMCS AND DEVICE-SPECIFIC CONTRAINDICATIONS: Transthoracic echocardiography (TTE) and—if indicated—transoesophageal echocardiography (TOE) offer important information on patient’s suitability for the different stMCS devices. Echocardiographers should be familiar with the absolute and relative contraindications of each stMCS device before implantation. In the absence of absolute contraindications, we should carefully look for risk indicators, which could interfere with optimum function of the stMCS devices10 (Table 1). IABP: Echocardiographic assessment is crucial to identify contraindications for intra-aortic balloon pump (IABP) (Figure 2). Moderate or greater aortic regurgitation (AR) can reduce the haemodynamic benefits of counterpulsation and exacerbate LV dilation and pulmonary congestion. Significant aortic atheroma increases the risk of balloon rupture, embolism, or atheroma dislodgment, necessitating TOE assessment of the aortic valve and thoracic aorta. IABP is ineffective in isolated RV failure and may worsen systemic hypotension due to insufficient LV preload.11,12 Additionally, it can aggravate vasoplegia or dynamic LVOT obstruction by reducing afterload. Anatomical contraindications, such as aortic dissection or coarctation, must also be ruled out. However, successful use in patients with Type A dissection has been reported following echocardiographic confirmation of placement in the true lumen.13 Therefore, while IABP is effective in select cases, careful echocardiographic evaluation is essential to optimize patient selection and avoid potential complications. LEFT-SIDED MICROAXIAL FLOW DEVICES (IMPELLA CP/5.5): The left-sided Impella (Figure 3) offers effective LV unloading and improved systemic perfusion; therefore, they are suitable for patients with isolated LV failure. Isolated or predominant RV failure should be ruled-out, as the left-sided Impella forward flow might aggravate the RV failure due to RV/Impella flow mismatch. In addition, in case of LV underfilling due to insufficient RV output, the left-sided Impella could precipitate LV chamber collapse, resulting in suction events and haemolysis. Traditional echocardiographic predictors of RV failure after initiation of left-sided durable continuous-flow support include a TAPSE < 7.5 mm and increased RV dimensions relative to the failing LV (i.e. RV to LV diameter ratio > 0.75). This rationale extends also to the risk stratification of RV failure after initiation of Impella support.14,15 In addition, a paradoxical leftward interventricular septum shifting suggests a prevailing RV dysfunction contributing to the clinical picture. To ensure adequate support, aortic valve competence is required, as significant AR may lessen antegrade blood flow and increase LV end-diastolic pressure due to blood recirculation between aorta and LV. Therefore, whenever more than moderate AR is observed, initiation of left-sided Impella should be considered on a case-by-case basis and weighted against other stMCS. Notably, Impella insertion has been safely and effectively performed in cases of both severe aortic stenosis, aortic bioprosthesis, and AR.16,17 Additional selected echocardiographic markers of higher risk for device-related complications include a small LV cavity and a mitro-aortic angle < 126.5°—all associated with a higher risk of suction and haemolysis.18 Conversely, a larger LV (LVEDD > 50.5 mm) conveys an increased risk of Impella device malrotation within the LV.19 Anatomical contraindications to Impella insertion include the presence of mechanical aortic prosthesis, aortic dissection, and LV mobile thrombi, all of which can be identified on echocardiography. RIGHT-SIDED STMCS DEVICES (PROTEKDUO/IMPELLA RP): Right-sided stMCS devices have somehow less availability compared with left-sided stMCS devices. When evaluating the use of RV support devices in patients with (isolated) RV failure, TTE and TOE play a crucial role in identifying potential contraindications or factors that may lead to device dysfunction. These include the presence of thrombi in the right atrium (RA) or RV, tricuspid or pulmonary valve stenosis, mechanical tricuspid or pulmonary prosthetic valves, or intracardiac shunts (such as atrial septal defect, ventricular septal defect, or patent foramen ovale) that could result in left-to-right shunting. Furthermore, echocardiography should evaluate pulmonary valve competence, as significant pulmonary regurgitation (PR) would decrease the flow to the pulmonary circulation and the RV unloading effect of the device. In addition, the presence of prominent, floating atrial structures (e.g. Chiari network and/or Eustachian valve) could interact with the cannula placement and interfere with the atrial drainage. Finally, echodynamic assessment should rule out severe pulmonary hypertension and signs of increased left atrial pressure as the first would diminish the forward flow generated by the ProtekDuo cannula at any given speed and the latter could precipitate pulmonary oedema as the antegrade flow of the device will worsen left chamber and pulmonary circulatory stasis. Tricuspid valve function should be assessed using 2D, color flow doppler, and spectral Doppler echocardiography.20 While TR is typically well tolerated during temporary right ventricular assist device support, the presence of more than mild PR may limit the efficacy of device outflow delivered to the PA.20 Of note, the presence of masses in the right heart chambers or PA could lead to device obstruction or pulmonary embolism.20 V-A ECMO: Echocardiography plays a crucial role in identifying or ruling out reversible cardiac pathologies that may be the primary cause of haemodynamic deterioration, such as cardiac tamponade, undiagnosed valvular disease, or LV dysfunction. It is also essential for detecting absolute contraindications to veno-arterial extracorporeal membrane oxygenation (V-A ECMO), such as aortic dissection. Significant AR is a relative contraindication to V-A ECMO, as the increased LV afterload can exacerbate AR, worsening LV distension and pulmonary congestion. Even moderate AR can cause significant LV distension and congestion at high V-A ECMO flows.21 Additionally, echocardiography provides valuable insights into aortic atherosclerosis, guiding the imaging specialist in selecting the optimal cannulation approach (central vs. peripheral) and technique (surgical vs. percutaneous) to minimize complications. Furthermore, it is instrumental in assessing right heart morphology, helping to identify structural abnormalities that could impede venous cannula placement, and ensuring effective device function and patient safety. Severe, spontaneous echo contrast is another high-risk marker for LV overload, stasis, and subsequent thrombosis.22 Finally, severe mitral regurgitation (MR) intrinsically carries a higher risk for pulmonary oedema and might be worsened by the retrograde V-A ECMO flow. IABP: The efficient operation of an IABP relies on proper positioning, with the balloon tip placed ∼2 cm distal to the origin of the left subclavian artery (LSA) and the proximal end positioned above the diaphragm (Figure 2). Echocardiographic assessment and guidance can be performed using TTE (off-axis apical two-chamber and suprasternal views)23 or TOE; however, TOE is often preferred for superior visualization of guidewire navigation in the aorta. Advancement of the femorally inserted guidewire can be monitored using short-axis and long-axis aortic views. The IABP tip should be positioned just below the origin of the LSA, which can be visualized by counterclockwise rotation of the TOE probe from an upper oesophageal aortic arch short-axis view, where the LSA appears at the 1 o’clock position. To position the IABP correctly, the echo-dense distal tip of the catheter and the lucent balloon body are distinguished from the thin guidewire. The origin of the LSA is identified, and the balloon tip is advanced to ∼2 cm below this point. A simple confirmation technique involves locating the balloon tip in a descending aorta short-axis view, placing the echocardiographer’s fingertips on the TOE probe at the patient’s teeth, and withdrawing the probe until the LSA origin appears. The distance from the fingers to the teeth represents the distance from the balloon tip to the artery, guiding final adjustments. If the LSA origin is not visible, the inferior margin of the aortic arch serves as an alternative landmark, with the IABP tip positioned just distal to this point. After final positioning, ensure there is no occlusion of the LSA or left carotid artery. LEFT-SIDED MICRO-AXIAL FLOW PUMP: Proper positioning of the guidewire and Impella in the LV should account for individual intracardiac anatomical variations (Figure 3). Accurate initial positioning is essential for optimal device function, preventing complications such as suction events and haemolysis, and facilitating later correction in case of inlet malrotation. Device insertion depends on the Impella type: Impella CP is typically implanted percutaneously via the femoral artery, while Impella 5.5 is inserted surgically via the axillary artery or direct aortic approach. Guidewire visualization in the aorta and exclusion of aortic dissection are best accomplished using TTE parasternal short- and long-axis aortic views for femoral access. Wire crossing of the aortic valve and intraventricular positioning of the wire, followed by the Impella, is best achieved using the mid-oesophageal long-axis view (or parasternal long-axis view in TTE). Misplacement of the wire across the mitral valve must be identified and corrected. The device inlet should be positioned freely within the LV cavity, directed towards the apex, without obstruction from the papillary muscle or mitral valve. After device placement, the guidewire is removed, and the final insertion depth is set. Recommended depths (measured from the device inlet to the aortic valve) are ∼3.5 cm for Impella CP and 4.5–5 cm for Impella 5.5, with slight adjustments based on individual anatomy to ensure optimal pump performance. The outlet area should be seen well above the aortic valve. The positions of the inlet and outlet zones can be inferred based on the artefacts they produce. In the mid-oesophageal long-axis view, the inlet zone, located in the LV far field, appears as a hypoechoic area due to acoustic shadowing. Meanwhile, the outlet area, seen in the near field within the ascending aorta, generates a prominent reverberation artefact.20,24 Additionally, when the device is correctly positioned, colour flow Doppler interrogation should reveal a mosaic pattern of turbulence around the outlet area, well above the aortic valve, confirming appropriate flow dynamics.20,24 In addition to device depth, proper orientation within the LV should always be ascertained. The device inlet should be positioned freely within the LV cavity, directed towards the apex. The abnormal orientation of Impella inlet and pig-tail away from the LV apex and towards the mitral valve inlet and the inferolateral LV wall has been termed ‘Impella malrotation’ and is associated with poor outcomes and worse LV unloading in CS patients.19,25 Echocardiography promptly detects malrotation and should systematically be done at insertion and at least daily during support. In addition, the presence of the ‘crushed pig-tail sign’ (i.e. an abrupt bend of the Impella pigtail) on chest X-ray can also suggest malrotation26 (Figure 3). PROTEKDUO/IMPELLA RP FLEX: TOE with colour Doppler is highly effective in evaluating the orientation of the outflow jet and the positioning of the outflow cannula (Figure 4). The cannula tip should be positioned 1–2 cm beyond the pulmonary valve within the main PA, ensuring that the flow is directed forward. The side holes at the distal end of the cannula facilitate optimal flow distribution into both the right and left PAs, promoting efficient circulation.27 For verification, several mid-oesophageal TOE views can be used, including the mid-oesophageal RV inflow outflow and upper oesophageal aortic arch short-axis views with colour Doppler and continuous-wave Doppler. IMPELLA RP: The device is positioned across the tricuspid and pulmonary valves, with the inflow port located in the IVC and the outflow port in the PA. Ideally, the outflow port should be positioned in the main PA, directed towards the left PA, allowing the Swan–Ganz catheter to be placed in the right PA.24 A bicaval view on TOE can help visualize the inflow port, typically positioned at the IVC/RA junction, while a mid-oesophageal RV inflow–outflow view and upper oesophageal views can confirm proper outflow port placement. Similar to left-sided Impella devices, the Impella right percutaneous (RP) is susceptible to migration, which may result in device malfunction. If the outflow port is positioned at or below the pulmonary valve, it can lead to reduced support or even complete recirculation within the RV. In such cases, TTE or TOE should be used to verify proper Impella positioning.28 V-A ECMO: After cannulating the femoral vein and/or artery, the guidewire’s passage into the IVC and/or aorta can be verified using TTE (mid-axillary transhepatic view) or TOE. The IVC is best visualized by advancing the probe from a modified mid-oesophageal bicaval view while the thoracic aorta can be assessed using the descending aorta long- and short-axis view.29 When inserting the guidewire via the right internal jugular vein, its passage through the superior vena cava (SVC) and into the RA can be confirmed in the mid-oesophageal bicaval view (Figure 5). Once proper guidewire placement is confirmed, percutaneous dilation and ECMO cannula insertion can be performed. The drainage cannula can then be advanced up the IVC into the RA, and femoral arterial cannula is not usually seen on TOE. MONITORING BEDSIDE DEVICE PERFORMANCE AND TROUBLESHOOTING IN THE INTENSIVE CARE UNIT: Multidisciplinary echo rounds might be valuable in daily bedside monitoring, assessing device performance and troubleshooting the following: ASSESSING LV UNLOADING AND DECOMPRESSION: V-A ECMO configuration poses the risk of increased LV afterload, potentially reducing LV ejection to the point of persistent, complete closure of the aortic valve.30 Serial assessment of LVOT VTI and visual assessment of aortic valve opening while increasing V-A ECMO flow help to identify worsening LV ejection and prompt venting strategies.31 Echocardiographic parameters of LV unloading in stMCS have been predominantly studied in durable continuous-flow LV assist devices.32 Key parameters with reasonable accuracy in distinguishing PCWP above or below 15 mmHg include right atrial pressure, systolic pulmonary artery pressure, left atrial volume index, E/A, and E/e′. Additionally, a relative change in LVEDd has been identified as a useful metric.31 However, individually, these parameters exhibit only moderate sensitivity and specificity. The presence of severe MR, lack of aortic valve opening, and spontaneous echo contrast in the LV are indicators of inadequate LV unloading. Data specific to stMCS remain limited, although ongoing studies on speckle tracking show promise pending further validation.33 DETECTING COMPLICATIONS: Different stMCS modalities exhibit varying complication rates.34 Daily focused echocardiographic assessments are advised to ensure correct device positioning and detect complications, including intracardiac or valvular thrombi and cardiac tamponade (Table 2). Haemodynamics of tamponade are modified because of ECMO and the effects of it may only become apparent when reducing the flow during a weaning trial.35 A pericardial haematoma that compresses cardiac chambers may not change haemodynamics if it does not affect cannula flow WEANING: stMCS weaning is the gradual reduction of the level of support to a minimum level with the goal of either de-escalation or explantation. However, before initiating the weaning process, it is essential to evaluate the patient’s readiness. ‘Readiness to wean’ indicates that the patient has demonstrated sufficient recovery across clinical, haemodynamic, metabolic, and imaging parameters to warrant a trial of reduced circulatory support, referred to as a weaning trial. Weaning from stMCS is primarily a clinical decision and does not always require full restoration of cardiac function. The most studied echocardiographic parameters for assessing readiness to wean are LVEF ≥ 20–25% and LVOT VTI ≥ 9–10 cm. However, these parameters are highly influenced by global haemodynamics, and observing an improving trend is often more critical than meeting absolute thresholds.35,36 Additionally, structural considerations play a key role in readiness to wean. Persistent severe MR or other structural abnormalities requiring intervention should be thoroughly evaluated before initiating the weaning process. A comprehensive echocardiographic assessment helps identify residual issues that could impact the success of support de-escalation or explantation (Figures 6 and 7). The ultimate goal is to predict stability after decannulation (ready to decannulate). This process varies depending on the specific device being used. Table 3 summarizes echocardiographic parameters used during weaning of MCS. FUTURE DIRECTIONS: Speckle-tracking echocardiography (STE) and 3D echocardiography have become well-established techniques, offering significant incremental value beyond conventional 2D imaging in assessing cardiac function, monitoring myocardial recovery, and guiding stMCS weaning. STE-derived myocardial strain is particularly useful for evaluating LV unloading, as demonstrated by Hammoudi et al.,33 who reported a strong linear relationship between STE strain and LV stroke work in post-myocardial infarction pigs undergoing acute LV unloading with Impella. Similarly, in a right coronary artery embolization-induced CS model, STE provided detailed insights into global and regional LV and RV function, distinguishing the effects of vasoactive therapy from mechanical support via Impella RP.37 In the context of stMCS weaning, myocardial recovery can be effectively monitored using LV and RV STE and 3D parameters, which have emerged as predictors of successful de-escalation.38,39 These techniques are also valuable for assessing myocardial viability, informing decisions on stMCS escalation, de-escalation, or transition to long-term support. While STE has been extensively used to evaluate myocardial viability in ischaemic heart disease, its application in stMCS remains an area for further exploration.40,41 Given the challenges of performing cardiac magnetic resonance in critically ill patients, STE could serve as a viable alternative for assessing myocardial function in this population. Another promising innovation is the integration of pressure-strain loops, which combine STE strain measurements with pressure data to assess myocardial work indices, providing non-invasive insights into ventricular-arterial-device coupling.42 Additionally, blood speckle imaging (BSI) has shown potential in stMCS management by directly tracking blood velocity vectors. BSI has been particularly useful in visualizing Impella inlets and identifying the watershed region in VA-ECMO, which represents the interface between antegrade flow from Impella and retrograde flow from ECMO. Continuous monitoring of the watershed region’s progression may offer valuable insights into readiness for ECMO weaning, making it an exciting avenue for future research.43 Beyond advanced functional imaging, ultrasound-enhancing agents (UEAs) have become integral in critically ill patients, particularly in post-cardiac surgery scenarios, where mediastinal changes, chest tubes, and mechanical ventilation can impair conventional echocardiography.44 The feasibility of UEA use in stMCS settings has been demonstrated in small studies, where it successfully identified the watershed area in the abdominal aorta in VA-ECMO patients and improved thrombus detection and LV function assessment without complications.45 However, safety concerns persist, particularly regarding the potential interference with arterial bubble sensors that may trigger automatic pump shutdowns. To mitigate these risks, the American Society of Echocardiography now recommends using UEAs in ECMO-supported patients only under the supervision of experienced perfusionists and clinicians.10 Artificial intelligence (AI) is transforming the management of stMCS by enabling real-time data analysis and supporting clinical decision-making. AI models can leverage complex patterns within patient data to identify risks and optimize treatment strategies throughout the stMCS continuum, including patient selection and risk stratification, device optimization and performance monitoring, and complication prediction and management. Machine learning (ML), encompassing supervised and unsupervised algorithms, has garnered significant attention in CS research.46 ML enables the analysis of large datasets, including patient demographics, clinical biomarkers, imaging data, and comorbidities, to identify CS phenotypes and predict outcomes. A particularly promising innovation is the concept of a digital twin, a virtual model of a real-life patient that receives real-time updates on physiological and treatment data. This ‘digital twin’ can simulate and personalize clinical outcomes using AI predictions, assisting clinicians in therapy regulation and long-term prognosis adjustments.47 Despite its potential, several challenges need to be addressed before AI can be fully realized in the echo management of stMCS. These include the need for standardized, high-quality data to train robust AI models, ensuring the safety and efficacy of algorithms, and addressing concerns around patient privacy, algorithm bias, equity in access, and the transparency of AI systems. CONCLUSIONS: Echocardiography plays a pivotal role in the selection, guidance, monitoring, and escalation/de-escalation of stMCS devices, providing real-time haemodynamic assessment, ensuring precise device placement, and detecting complications. In the management of patients with stMCS devices, ultrasound plays a pivotal role beyond its established applications. Specific
PMID:40495936 | 1 | 40495936 | PMC12150027 | 10.1093/ehjimp/qyaf067 | fetched |  | ally, ultrasound guidance enhances the safety and success rate of vascular access procedures, which are critical in the placement and management of MCS devices. By providing real-time visualization of vascular anatomy, ultrasound reduces the risk of complications such as arterial puncture, haematoma, and catheter malposition, thereby improving patient outcomes. Furthermore, ultrasound serves as a non-invasive tool to assess pulmonary congestion and systemic perfusion in patients supported by MCS devices. Lung ultrasound, for instance, can detect B-lines indicative of pulmonary oedema, allowing for timely interventions to manage fluid status and prevent respiratory complications. Additionally, Doppler ultrasound can evaluate systemic perfusion by assessing blood flow in various organs, aiding in the optimization of MCS device settings and overall patient management.48 However, several challenges and pitfalls impact its utility in critically ill patients. However, its implementation requires systematic imaging protocols, standardized interpretation and storage, and continuous education to maintain quality and clinical utility. The acquisition of Doppler-derived and myocardial performance indices is often challenging in patients with mechanical ventilation, post-cardiotomy status, anasarca, valvular disease, or atrial rhythm disorders, potentially limiting accurate assessment. Additionally, real-time quality assurance and imaging consistency remain concerns, particularly for handheld ultrasound devices, reinforcing the importance of interdisciplinary imaging expertise in critical care settings. Given the complexity of stMCS management, integrating an imaging specialist into the multidisciplinary team is crucial to ensuring accurate assessments, guiding therapy adjustments, and optimizing clinical outcomes. Establishing clear diagnostic and monitoring protocols will improve clinical consistency, while validating advanced techniques, such as strain imaging and pressure-strain loops, could enhance patient selection for device de-escalation or transition to long-term support. In conclusion, while echocardiography remains indispensable in stMCS management, ongoing research, validation, and standardization are essential to optimizing its role, refining patient selection, and improving clinical outcomes through AI, multimodality imaging, and echodynamic advancements.
PMID:24139614 | 0 | 24139614 | PMC3988268 | 10.1016/j.jtcvs.2013.08.044 | fetched | {"title": "End-diastolic flow reversal limits the efficacy of pediatric intra-aortic balloon pump counterpulsation.", "authors": ["Bartoli CR", "Rogers BD", "Ionan CE", "Pantalos GM"], "journal_abbrev": "J Thorac Cardiovasc Surg", "year": 2014, "volume": "147", "issue": "5", "pages": "1660-7", "doi": "10.1016/j.jtcvs.2013.08.044", "pmid": "24139614", "pmcid": "PMC3988268", "pub_types": ["Comparative Study", "Journal Article", "Research Support, N.I.H., Extramural", "Research Support, Non-U.S. Gov't"]} | End-diastolic flow reversal limits the efficacy of pediatric intra-aortic balloon pump counterpulsation. BACKGROUND: Counterpulsation with an intra-aortic balloon pump (IABP) has not achieved the same success or clinical use in pediatric patients as in adults. In a pediatric animal model, IABP efficacy was investigated to determine whether IABP timing with a high-fidelity blood pressure signal may improve counterpulsation therapy versus a low-fidelity signal. METHODS: In Yorkshire piglets (n = 19; weight, 13.0 +/- 0.5 kg) with coronary ligation-induced acute ischemic left ventricular failure, pediatric IABPs (5 or 7 mL) were placed in the descending thoracic aorta. Inflation and deflation were timed with traditional criteria from low-fidelity (fluid-filled) and high-fidelity (micromanometer) blood pressure signals during 1:1 support. Aortic, carotid, and coronary hemodynamics were measured with pressure and flow transducers. Myocardial oxygen consumption was calculated from coronary sinus and arterial blood samples. Left ventricular myocardial blood flow and end-organ blood flow were measured with microspheres. RESULTS: Despite significant suprasystolic diastolic augmentation and afterload reduction at heart rates of 105 +/- 3 beats per minute, left ventricular myocardial blood flow, myocardial oxygen consumption, the myocardial oxygen supply/demand relationship, cardiac output, and end-organ blood flow did not change. Statistically significant end-diastolic coronary, carotid, and aortic flow reversal occurred with IABP deflation. Inflation and deflation timed with a high-fidelity versus low-fidelity signal did not attenuate systemic flow reversal or improve the myocardial oxygen supply/demand relationship. CONCLUSIONS: Systemic end-diastolic flow reversal limited counterpulsation efficacy in a pediatric model of acute left ventricular failure. Adjustment of IABP inflation and deflation timing with traditional criteria and a high-fidelity blood pressure waveform did not improve IABP efficacy or attenuate flow reversal. End-diastolic flow reversal may limit the efficacy of IABP counterpulsation therapy in pediatric patients with traditional timing criteria. Investigation of alternative deflation timing strategies is warranted. ABSTRACT: OBJECTIVECounterpulsation with an intraaortic balloon pump (IABP) has not achieved the same successes or clinical use in pediatric patients as in adults. In a pediatric animal model, IABP efficacy was investigated to determine whether IABP timing with a high-fidelity blood pressure signal may improve counterpulsation therapy versus a low-fidelity signal.METHODSIn Yorkshire piglets (n=19, 13.0±0.5 kg) with coronary ligation-induced acute ischemic left ventricular failure, pediatric IABPs (5 or 7cc) were placed in the descending thoracic aorta. Inflation and deflation were timed with traditional criteria from low-fidelity (fluid-filled) and high-fidelity (micromanometer) blood pressure signals during 1:1 support. Aortic, carotid, and coronary hemodynamics were measured with pressure and flow transducers. Myocardial oxygen consumption was calculated from coronary sinus and arterial blood samples. Left ventricular myocardial blood flow and end-organ blood flow were measured with microspheres.RESULTSDespite significant suprasystolic diastolic augmentation and afterload reduction at heart rates of 105±3bmp, left ventricular myocardial blood flow, myocardial oxygen consumption, the myocardial oxygen supply/demand relationship, cardiac output, and end-organ blood flow did not change. Statistically significant end-diastolic coronary, carotid, and aortic flow reversal occurred with IABP deflation. Inflation and deflation timed with a high-fidelity versus low-fidelity signal did not attenuate systemic flow reversal or improve the myocardial oxygen supply/demand relationship.CONCLUSIONSSystemic end-diastolic flow reversal limited counterpulsation efficacy in a pediatric model of acute left ventricular failure. Adjustment of IABP inflation and deflation timing with traditional criteria and a high-fidelity blood pressure waveform did not improve IABP efficacy or attenuate flow reversal. End-diastolic flow reversal may limit the efficacy of IABP counterpulsation therapy in pediatric patients with traditional timing criteria. Investigation of alternative deflation timing strategies is warranted. INTRODUCTION: Counterpulsation with an intraaortic balloon pump (IABP) is the most common mechanical circulatory support strategy for a variety of cardiac pathologies in adults1. Yet IABP therapy in pediatric patients has not demonstrated the same degree of efficacy and has not gained widespread clinical use2. Differences between adult and pediatric anatomy and physiology may limit the efficacy of IABP therapy in neonates, infants, and children. In this study, we examined IABP counterpulsation in a pediatric model of acute ischemic left ventricular failure. We tested the hypothesis that the efficacy of pediatric IABP therapy is improved with a high-fidelity (micromanometer) rather than a traditional low-fidelity (fluid-filled) arterial blood pressure signal used to adjust IABP inflation and deflation timing. METHODS: All animals received humane care and were handled in accordance with the Guide for the Care and Use of Laboratory Animals (National Research Council, 1996). Experimental procedures followed animal study protocols approved by the University of Louisville Institutional Animal Care and Use Committee. EXPERIMENTAL DESIGN: Piglets (n=19, 13.0±0.5 kg) were instrumented surgically to determine aortic, carotid, and coronary artery hemodynamics. Myocardial oxygen consumption (MVO2) was calculated from coronary sinus and arterial blood gas data. Left ventricular myocardial blood flow and regional end-organ blood flow were determined with neutron-activated 15 μm microspheres. Sequential coronary ligation was performed to induce acute ischemic left ventricular failure. A pediatric IABP was placed in the descending thoracic aorta. In each animal, hemodynamic waveforms (15 second data epochs), blood gasses, and end-organ blood flows were measured at steady-state during the following experimental test conditions: 1) normal baseline (IABP off), 2) coronary ligation-induced left ventricular failure (IABP off), 3) IABP support with timing adjusted to the arterial blood pressure waveform acquired with a low-fidelity (<20 Hz response) fluid-filled catheter in the radial artery, and 4) IABP support with timing adjusted to the arterial pressure waveform acquired with a high-fidelity (<5 kHz repsonse) micromanometer pressure transducer in the aortic root. Measurements were recorded during 1:1 counterpulsation support in which each aortic valve closure initiated rapid balloon inflation, and each end-diastole initiated rapid balloon deflation. An aortic root pressure waveform transduced by a high-fidelity catheter results in a more precise waveform morphology with a more pronounced dicrotic notch, less signal gain, and less phase distortion3. Consequently, a high-fidelity signal was presumed to enable more accurate inflation and deflation timing of the IABP. SURGICAL PREPARATION: Animals were fasted, preanesthetized with intramuscular ketamine (30 mg/kg) and acepromazine (0.2 mg/kg), and anesthetized with isoflurane (1.5 to 3%) and room air. A left lateral thoracotomy was performed and the left and right carotid artery and jugular vein were exposed through neck incisions. A single-tip high-fidelity micromanometer catheter (2.5 Fr, SPR-524, Millar Instruments, Houston, TX) was placed in the left atrium. A dual-tip high-fidelity micromanometer catheter (5 Fr, SPC-721, Millar Instruments, Houston, TX) was advanced retrograde from the ascending aorta into the left ventricle. Transit-time ultrasonic flow probes (T205, Transonics, Ithaca, NY) were placed around the aortic root, left carotid artery, and left anterior descending (LAD) coronary artery to measure volumetric blood flows. Sampling catheters were introduced into the coronary sinus via the hemiazygous vein, pulmonary artery, and right carotid artery to sample blood and measure blood gases. An infusion catheter was introduced into the left atrium for serial injections of microspheres during each experimental test condition to determine end-organ blood flows throughout the body4. A blood-sampling catheter was introduced into the right femoral artery for simultaneous withdrawal of a microsphere reference blood-flow sample during each test condition. A fluid-filled pressure-monitoring catheter (24 gauge Angiocath, Becton Dickinson, Sandy, UT) was inserted into the right radial artery. The animal was anticoagulated with a bolus injection of heparin (300 units/kg) and subsequent small boluses of heparin to maintain an ACT greater than 300 seconds. An IABP catheter (5 or 7 cc, Datascope, Fairfield, NJ) was inserted via a left femoral artery cut-down and advanced until the tip was distal to the left subclavian artery as confirmed by digital palpation. Balloon timing was manually adjusted with the arterial pressure waveform (fluid-filled radial artery catheter or aortic root Millar catheter) displayed on the IABP console (System 97, Datascope, Fairfield, NJ). Balloon inflation and deflation were timed to maximize aortic diastolic pressure augmentation and minimize aortic end-diastolic pressure (afterload). INDUCTION OF ISCHEMIC LEFT VENTRICULAR FAILURE: Intravenous lidocaine (20 mg bolus, 20 mg/hour infusion) and esmolol (5 mg/kg bolus, 50 μg/kg/min infusion) were administered to prevent arrhythmia. Sequential coronary ligation of branches of the LAD induced acute ischemic left ventricular failure. Target cardiac dysfunction was achieved when three of four criteria were met: approximate reduction of 1) left ventricular cardiac output by 25%, 2) mean aortic pressure by 10 mmHg, 3) mixed venous O2 saturation by 10%, and 4) an elevation of left atrial pressure (LAP) and/or left ventricular end-diastolic pressure by 5 mmHg. Animals with life-threatening hypotension were supported with boluses of intravenous normal saline and continuous infusion of phenylephrine and/or epinephrine to effect. BLOOD GAS ANALYSIS: During each experimental test condition, blood samples were simultaneously withdrawn from the coronary sinus, pulmonary artery, and right carotid artery. Samples were processed with a blood gas analyzer (IRMASL Blood Analysis System, Diametrics Medical, Roseville, MN) to measure hemoglobin content ([Hb] g/dL) and hemoglobin oxygen saturation (% O2 Sat). Oxygen content ([O2]) was calculated for each blood sample: [O2]=%O2Sat×[Hb]×O2capacityofHb(1.34mlO2/g)100=[mlO2/100ml] Total myocardial oxygen consumption (MVO2) was calculated as: MVO2=([O2]a)-([O2]cs)×MBF=(mlO2/min/100g) END-ORGAN BLOOD FLOW MEASUREMENTS: During each experimental test condition, a different color (isotope label) of 15 μm neutron-activated microspheres (1×106 Microspheres in 0.4 ml suspension, Biopal, Worcester, MA) was injected into the left atrium followed by a 4 ml saline flush. The microsphere technique enabled the measurement of regional end-organ blood flow in vascular beds of interest as previously described4,5. During microsphere injection, a reference blood-flow sample was drawn from the femoral artery at a rate of 4 ml/min for 60 seconds with a calibrated syringe pump (Harvard Apparatus, Holliston, MA). The withdrawal sample acted as a reference to determine organ specific flows in ml/min/g of tissue4. NECROPSY AND MICROSPHERE ANALYSIS: After completion of the experimental protocol, animals were euthanized with an increase in anesthetic depth and an intravenous bolus injection of supersaturated KCl. After euthanasia, the heart was harvested and weighed. The ventricles were dissected from the atria. The left ventricle was dissected form the right ventricle. The brain, lungs, kidneys, pancreas, liver, spleen, and adrenals were harvested and weighed. Tissue and reference blood samples were sent to BioPAL, Inc. (Worcester, MA) for radioactive assay and automated calculation of blood flow in ml/min/g for each sample during each experimental test condition. Tissue and blood samples were bombarded with neutrons to transiently activate each group of microspheres with a separate isotope label. The level of radioactivity detected in the sample was directly proportional to the number of microspheres present in the sample. The activity of microspheres in the reference blood sample was compared to the activity of microspheres that lodged in a tissue sample of interest. The ratio between the two activity counts was equal to the ratio between the calibrated rate of aortic withdrawal (known, 4 ml/min) and flow in the tissue of interest (unknown). INSTRUMENTATION, DATA REDUCTION, AND STATISTICS: All transducers were pre- and post-calibrated against known physical standards to ensure measurement accuracy. Collected data were signal conditioned (1000X gain and 60 Hz low-pass filter) and analog-to-digital converted at a sampling rate of 400 Hz for digital analysis using our GLP compliant data acquisition system6. Hemodynamic parameters were calculated on a beat-to-beat basis for each 15 second data set with the Hemodynamic Evaluation and Assessment Research Tool (HEART) program7 developed in Matlab (Version 6.5, MathWorks, Natick, MA). All analyzed beats in each data set (approximately 25 to 30 beats/15 second data set) were averaged to obtain a single representative mean value for each calculated parameter. Pressure and flow waveforms were used to derive the following hemodynamic parameters: heart rate (HR), aortic pressure (AoP), left atrial pressure (LAP), left ventricular pressure (LVP), aortic flow (AoF), carotid artery flow (CarotidF), and LAD flow (LADF). Left ventricular myocardial blood flow (LVF, as determined by microspheres) normalized to MVO2 was used as an index of the oxygen supply/demand relationship. Prism (version 4.00, GraphPad, La Jolla, CA) was used to perform statistical analyses and to plot data. To verify the induction of a clinically relevant state of left ventricular failure, paired student t-tests were used to compare values for LADF, AoF, AoP, mixed venous O2 saturation, end-organ-blood flows, LAP, and LVP at baseline and after coronary ligation. One-way repeated measures ANOVA with Tukey post-test for comparison of means were performed to compare left ventricular failure and IABP timing modes (low-fidelity and high-fidelity) for each parameter. A p<0.05 (95% confidence) was considered statistically significant: * vs. baseline; † vs. left ventricular failure. All data were presented as mean ± standard error. MODEL OF PEDIATRIC LEFT VENTRICULAR FAILURE: Ten of the 19 piglets completed the experimental protocol. In these animals, LAD ligation produced ischemic left ventricular failure with phenotypic similarities to clinical left ventricular failure (Table 1). Animals exhibited an approximate reduction in LADFmean by 40% (26±4 to 16±3 ml/min, p<0.01), AoF by 30% (2.4±2 to 1.7±0.2 L/min, p<0.001), AoPmean by 15 mmHg (67±4 to 53±3 mmHg, p<0.01), mixed venous O2 saturation by 15% (86±1 to 70±4%, p<0.001), renal blood flow by 45–50% (left kidney: 1.69±0.29 to 0.90±0.13 ml/min/100g, p<0.05; right kidney: 2.13±0.45 to 1.02±0.18 ml/min/100g, p<0.05), and an increased LAP by approximately 3 mmHg (9±1 to 12±1, p<0.001) and LVPend-diastolic by approximately 5 mmHg (9±1 to 14±1, p<0.01). The other nine animals died prior to completion of the study protocol, typically from intractable arrhythmia. Data from these animals were not included in any of the analyses. DIASTOLIC AUGMENTATION AND LEFT VENTRICULAR AFTERLOAD REDUCTION: At heart rates of 105±3 bpm (range 83 to 125 bpm), it was possible to achieve statistically significant suprasystolic diastolic pressure augmentation during each native cardiac beat (Figure 1 A, bottom waveform). During 1:1 IABP support timed with a low-fidelity (fluid-filled) or a high-fidelity (micromanometer) blood pressure signal, similar aortic pressure augmentation of greater than 20 mmHg was observed (Figure 1 B, p<0.0001). Significant left ventricular afterload reduction, as indicated by aortic end-diastolic pressure, was also achieved (Figure 1 A, middle and bottom waveforms; Figure 2 A). During 1:1 IABP support, aortic end-diastolic pressure decreased by approximately 8 mmHg (left ventricular failure, 41±2 mmHg; low-fidelity signal, 34±3 mmHg, p<0.0001; high-fidelity signal, 33±3 mmHg, p<0.0001). The increase in left ventricular myocardial blood flow during IABP support was small and did not reach statistical significance with either a low-fidelity or high-fidelity blood pressure signal (Figure 2 B). Similarly, myocardial oxygen consumption did not change during 1:1 IABP support with either a low-fidelity or high-fidelity blood pressure signal (Figure 2 C). Despite suprasystolic diastolic augmentation and left ventricular afterload reduction, the myocardial oxygen supply/demand relationship did not significantly improve with either a low-fidelity or a high-fidelity signal to adjust IABP inflation and deflation timing (Figure 2 D). SYSTEMIC FLOW REVERSAL: Brief but statistically significant end-diastolic flow reversal occurred in the aorta, carotid artery, and LAD (Figure 3). During 1:1 IABP support, mean flow reversal in the aorta increased by approximately 40 ml/min (left ventricular failure, −65±10 ml/min; low-fidelity signal, −97±12 ml/min, p<0.0001; high-fidelity signal, −111±12 ml/min, p<0.0001). Mean flow reversal in the carotid artery increased by approximately 15 ml/min (left ventricular failure, −3±2 ml/min; low-fidelity signal, −19±3 ml/min, p<0.0001; high-fidelity signal, −20±3 ml/min, p<0.0001). Mean flow reversal in the LAD increased by approximately 0.70 ml/min (left ventricular failure, −0.24±0.06 ml/min; low-fidelity signal, −1.01±0.24 ml/min, p<0.0001; high-fidelity signal, −0.91±0.18 ml/min, p<0.0001). END-ORGAN BLOOD FLOW: Cardiac output (AoF) did not significantly improve during 1:1 IABP support (left ventricular failure 1.67±0.15 L/min; low-fidelity signal, 1.78±0.15 L/min; high-fidelity signal, 1.84±0.15 L/min). As a result, systemic end-organ blood flow did not significantly improve with 1:1 IABP support timed with a low-fidelity or a high-fidelity blood pressure signal (Table 2). DISCUSSION: Pediatric patients with life-threatening heart failure present difficult clinical challenges. Over the past five decades, the intraaortic balloon pump (IABP) has successfully supported pediatric patients with congenital and acquired heart disease as a bridge to decision, bridge to cardiac transplantation, and bridge to recovery. Yet mortality rates greater than 35%2,8,9 and the lowest long-term survival for any pediatric mechanical circulatory support modality10 indicate a need to improve the efficacy of pediatric IABP therapy. In this study, in a piglet model of acute left ventricular failure, 1:1 IABP counterpulsation was possible at high heart rates. However, brief systemic blood flow reversal coincident with IABP deflation may have limited the efficacy of IABP support. Despite significant suprasystolic diastolic augmentation and significant left ventricular afterload reduction during 1:1 IABP support, left ventricular myocardial blood flow, myocardial oxygen consumption, and the myocardial oxygen supply/demand relationship did not improve. Cardiac output and end-organ blood flow also did not improve significantly. A brief but significant flow reversal in conduit arteries may have limited systemic forward flow. Traditional IABP inflation and deflation timing with a high-fidelity trigger did not attenuate flow reversal or improve IABP efficacy. These findings suggest that end-diastolic flow reversal is an important and novel mechanism that may limit the efficacy of IABP counterpulsation in pediatric patients with current inflation and deflation criteria. CLINICAL CONSIDERATIONS: The optimal timing of pediatric IABP inflation and deflation is uncertain11,12. In any patient with an IABP, it is traditionally thought that inflation and deflation of the balloon must occur during diastole between closure of the aortic valve and the subsequent systole. In an adult patient, this timing paradigm typically is not difficult to achieve. However, pediatric patients have higher heart rates than adults (newborn, 120–160 bpm; infant/toddler, 90–130 bpm), and diastole is short. As heart rate increases, diastole becomes disproportionally shorter13. As a result, inflation and deflation of the balloon must occur within a smaller window of time. From the current study, we speculate that the short duration of diastole during high heart rates may limit the hemodynamic benefits of IABP inflation. Correspondingly, rapid balloon deflation produces a sharp reduction in aortic end-diastolic blood pressure and brief flow reversal sufficient to limit net augmentation of coronary blood flow or improve the myocardial oxygen supply/demand relationship. Indeed, in adult patients, balloon deflation may produce a steep reduction of central end-diastolic blood pressure that may lead to a “steal” phenomenon during which coronary14,15, carotid15,16, cerebral17, and aortic15,18 blood flow reverse. Flow reversal is brief, and net flow augmentation is still positive. However, in a pediatric patient with a high heart rate, forward flow from diastolic augmentation occurs over a shorter period of time, and rapid balloon deflation occurs more frequently. As a result, the ratio between flow reversal and augmented forward flow is greater and may limit maximum net forward flow. Importantly, coronary flow reversal may be attenuated by adjusting inflation and deflation timing14. In adults with a normal heart rate (60 to 80 bpm), sufficient time exists between each systole to allow a brief (up to 25 ms) deflation delay. As a result, subtle variations in the timing of inflation and deflation may permit an incremental tradeoff between afterload and myocardial workload reduction versus diastolic augmentation and coronary perfusion. Unfortunately, in pediatric patients, high heart rates with a shorter diastole minimize the available window to delay deflation for this tradeoff. Similarly, brief timing errors associated with a fluid-filled versus a high-fidelity signal12 or high-fidelity signal versus echocardiographic timing19 may reduce counterpulsation efficacy in pediatric patients. A precise signal is required to accurately time inflation and deflation and maximize counterpulsation efficacy in pediatric patients. To this end, integration of a high-fidelity micromanometer into the catheter tip may eliminate timing error(s), especially compared to low-fidelity fluid-filled catheters, which introduce gain and phase distortion and tend to drift over time20. Indeed, we and other investigators19 observed that it was easier to identify inflation and deflation timing landmarks in the arterial pressure waveform acquired with a high-fidelity signal. Consequently, timing errors were unlikely responsible for end-diastolic flow reversal. In our study, adjustment of IABP timing with a high-fidelity blood pressure signal (that minimized timing error) did not significantly reduce end-diastolic flow reversal or improve hemodynamics or end-organ blood flow compared to a low-fidelity signal. This finding further supports the notion that balloon inflation and deflation with traditional timing criteria triggered by a high-fidelity arterial blood pressure signal is insufficient to improve counterpulsation therapy in pediatric patients. Instead, the development of alternative timing strategies may be needed to prevent end-diastolic flow reversal and improve IABP efficacy in pediatric patients. OTHER CONSIDERATIONS: Counterpulsation therapy is most effective when the displaced blood volume of the balloon and the left ventricular stroke volume are similar21,22. In this study we used 5 and 7 cc IABPs, which are the clinically recommended sizes for an 8 to 18 kg infant23. However, the stroke volume of a 12 kg piglet (and 12 kg human) with left ventricular failure is approximately 15 to 17 ml, more than three times the displacement volume of these IABPs. This observation begs the question of whether a larger balloon (and larger displaced blood volume) may produce greater diastolic pressure augmentation to significantly improve diastolic coronary blood flow. By the same token, rapid deflation of a balloon with a larger volume may produce a greater drop in end-diastolic aortic blood pressure and further increase the volume of blood flow reversal. This hypothesis remains untested. However, from a technical standpoint, the size of peripheral arteries in pediatric patients may limit the size of the balloon catheter that may be safely inserted and preclude counterpulsation therapy with a larger-sized IABP. Arterial compliance also plays an important role in counterpulsation therapy. It has been demonstrated in vitro24, experimentally in animal studies25, and clinically26 that the efficacy of counterpulsation therapy correlates inversely with arterial compliance. Specifically, stiffer arteries generate the greatest diastolic augmentation and reduction in end-diastolic blood pressure. As a result, it has been speculated that the high compliance of the pediatric aorta may limit IABP efficacy27. This effect has not been rigorously defined. Moreover, there is evidence to the contrary that the compliance of the pediatric aorta28 is sufficient to produce diastolic augmentation, afterload reduction, and improved hemodynamics27. Our results further support this point. FUTURE INVESTIGATIONS: Ongoing studies will determine whether alternative timing strategies such as delayed balloon deflation may improve the efficacy of pediatric counterpulsation. Similarly, the use of larger pediatric balloons with greater displacement volumes that are more comparable to native stroke volume may improve the efficacy of pediatric IABP counterpulsation. CONCLUSIONS: An incomplete understanding of couterpulsation has limited the widespread application of IABPs in pediatric patients. In a piglet model of acute ischemic left ventricular failure, significant diastolic augmentation and afterload reduction were achieved with a pediatric IABP at high heart rates. However, significant aortic, carotid, and coronary blood flow reversal coincident with balloon deflation during end-diastole reduced net systemic forward flow. IABP inflation and deflation timed with traditional criteria and a high-fidelity blood pressure signal did not attenuate flow reversal or improve IABP efficacy. These findings suggest that brief end-diastolic flow reversal is an important and novel mechanism that may limit the effectiveness of IABP counterpulsation in pediatric patients with current inflation and deflation criteria.
PMID:1598882 | 0 | 1598882 |  | 10.1016/0002-9149(92)90716-c | fetched | {"title": "Observations of coronary flow augmentation and balloon function during intraaortic balloon counterpulsation using transesophageal echocardiography.", "authors": ["Katz ES", "Tunick PA", "Kronzon I"], "journal_abbrev": "Am J Cardiol", "year": 1992, "volume": "69", "issue": "19", "pages": "1635-9", "doi": "10.1016/0002-9149(92)90716-c", "pmid": "1598882", "pmcid": null, "pub_types": ["Journal Article"]} | Observations of coronary flow augmentation and balloon function during intraaortic balloon counterpulsation using transesophageal echocardiography. The intraaortic balloon pump has been shown to decrease myocardial oxygen demand by afterload reduction, while increasing myocardial oxygen supply by diastolic augmentation of coronary blood flow. This diastolic augmentation of coronary flow has been demonstrated experimentally with invasive methods. Noninvasively, transesophageal echocardiography has demonstrated efficacy in enabling visualization of the proximal left coronary artery and in recording coronary blood flow velocity. To assess the potential of this technique in demonstrating quantitatively the increase in coronary flow during counterpulsation, 6 patients were studied during intermittent balloon pumping. Peak diastolic coronary blood flow velocity increased by a mean of 117% (range 62 to 287) during balloon inflation (p = 0.002). Furthermore, coronary flow velocity integral increased by a mean of 87% (range 43 to 176; p = 0.003). Problems associated with intraaortic balloon pumping were discovered by transesophageal echocardiography in 4 patients (incorrect balloon placement, damage to the aortic wall [2 patients], and premature balloon deflation time). Transesophageal echocardiography can be used in evaluating intraaortic balloon positioning within the aorta and in monitoring coronary artery flow augmentation during counterpulsation. This relatively noninvasive technique adds another dimension to the evaluation of balloon function and may help in optimizing the benefits of counterpulsation.
PMID:41348153 | 0 | 41348153 | PMC12680806 | 10.1186/s13089-025-00460-2 | fetched | {"title": "Optimizing IABP-patient interaction in VA-ECMO via transcranial doppler.", "authors": ["Yepes-Velasco AF", "Moreno-Araque JL", "Garzon-Posada N"], "journal_abbrev": "Ultrasound J", "year": 2025, "volume": "17", "issue": "1", "pages": "65", "doi": "10.1186/s13089-025-00460-2", "pmid": "41348153", "pmcid": "PMC12680806", "pub_types": ["Journal Article"]} | Optimizing IABP-patient interaction in VA-ECMO via transcranial doppler. BACKGROUND: Patients supported with veno-arterial extracorporeal membrane oxygenation may receive an intra-aortic balloon pump to reduce left ventricular afterload and improve aortic diastolic pressure. However, the effect of this combined mechanical support on cerebral hemodynamics is not uniform and can be influenced by intra-aortic balloon pump timing. Bedside transcranial Doppler offers a rapid, noninvasive way to detect maladaptive cerebral flow patterns and to guide patient-device interaction in real time. CASE PRESENTATION: We describe a postcardiotomy adult patient on peripheral veno-arterial extracorporeal membrane oxygenation with concomitant intra-aortic balloon pump assistance (1:1) who developed a reduction in cerebral oximetry. Transcranial Doppler of the middle cerebral artery showed increased pulsatility and reduced diastolic velocity, findings consistent with a transient decrease in cerebral perfusion pressure and compatible with balloon deflation asynchrony. Temporary suspension of balloon assistance improved the waveform. Deflation was then synchronized with the electrocardiogram so that it was completed at the onset of systole. Repeat transcranial Doppler performed minutes later showed restoration of diastolic flow and a lower pulsatility index, while extracorporeal support was maintained unchanged and the patient remained hemodynamically stable. CONCLUSIONS: In patients receiving veno-arterial extracorporeal membrane oxygenation and intra-aortic balloon pump support, cerebral blood flow may deteriorate if balloon timing is not aligned with the native cardiac cycle. Transcranial Doppler can detect these timing-related neurohemodynamic alterations at the bedside and can confirm their reversibility after simple, ECG-guided optimization of deflation. Integrating transcranial Doppler into multiparametric monitoring may help personalize mechanical circulatory support and protect cerebral perfusion in this high-risk population. ABSTRACT: BackgroundPatients supported with veno-arterial extracorporeal membrane oxygenation may receive an intra-aortic balloon pump to reduce left ventricular afterload and improve aortic diastolic pressure. However, the effect of this combined mechanical support on cerebral hemodynamics is not uniform and can be influenced by intra-aortic balloon pump timing. Bedside transcranial Doppler offers a rapid, noninvasive way to detect maladaptive cerebral flow patterns and to guide patient–device interaction in real time.Case presentationWe describe a postcardiotomy adult patient on peripheral veno-arterial extracorporeal membrane oxygenation with concomitant intra-aortic balloon pump assistance (1:1) who developed a reduction in cerebral oximetry. Transcranial Doppler of the middle cerebral artery showed increased pulsatility and reduced diastolic velocity, findings consistent with a transient decrease in cerebral perfusion pressure and compatible with balloon deflation asynchrony. Temporary suspension of balloon assistance improved the waveform. Deflation was then synchronized with the electrocardiogram so that it was completed at the onset of systole. Repeat transcranial Doppler performed minutes later showed restoration of diastolic flow and a lower pulsatility index, while extracorporeal support was maintained unchanged and the patient remained hemodynamically stable.ConclusionsIn patients receiving veno-arterial extracorporeal membrane oxygenation and intra-aortic balloon pump support, cerebral blood flow may deteriorate if balloon timing is not aligned with the native cardiac cycle. Transcranial Doppler can detect these timing-related neurohemodynamic alterations at the bedside and can confirm their reversibility after simple, ECG-guided optimization of deflation. Integrating transcranial Doppler into multiparametric monitoring may help personalize mechanical circulatory support and protect cerebral perfusion in this high-risk population. BACKGROUND: Veno-arterial extracorporeal membrane oxygenation (VA-ECMO) is a temporary circulatory mechanical support used for severe cardiogenic shock or cardiac arrest to restore systemic perfusion while gas exchange occurs extracorporeally [1, 2]. Because VA-ECMO does not fully unload the left ventricle, adjunct strategies are often required. In this setting, the intra-aortic balloon pump (IABP) is frequently used to reduce effective arterial elastance and left-ventricular afterload and to augment diastolic aortic pressure [3, 4]. Reports on cerebral hemodynamics during combined VA-ECMO + IABP are heterogeneous, reflecting variability in residual native ejection and device timing [5, 6]. Point-of-care neurosonography (TCD/TCCD) offers a practical, non-invasive way to trend cerebral hemodynamics at the bedside and to contextualize device interactions in patients on VA-ECMO with IABP [6–9]. We present a postcardiotomy patient with a transient reduction in cerebral perfusion detected by TCD and corrected through TCD-guided optimization of IABP timing, and we outline a focused, sequential monitoring approach for similar cases. CASE PRESENTATION: A 62-year-old patient with a mechanical aortic valve replacement (1985) was admitted with congestive heart failure due to atrial flutter and severe mitral insufficiency. After medical stabilization, the patient underwent mechanical mitral valve replacement with a St. Jude Regent #33 prosthesis, combined with anomalous pathway ablation, maze surgery (cryo and radiofrequency, biatrial), left atrial appendage exclusion, and intracavitary thrombus resection. The procedure had an ischemia time of 147 min and perfusion time of 190 min. Postoperatively, the patient developed severe oxygenation impairment and refractory hemodynamic deterioration, requiring peripheral veno-arterial femoro-femoral ECMO and intra-aortic balloon pump (IABP) support. During bedside monitoring with the IABP synchronized at 1:1 and ECMO flow at 4.2 L, a reduction in bilateral cerebral oximetry (NIRS) was observed. Transcranial Doppler revealed a low-flow, high-resistance cerebral pattern (Fig. 1), and the echocardiogram showed severely reduced systolic function LVEF < 30%. When IABP assistance was temporarily paused—without changing ECMO flow or fresh gas settings—immediate improvements in pulsatility indices and flow were observed (Fig. 1). These findings suggested a transient cerebral hypoperfusion pattern (characterized by increased pulsatility index with reduced diastolic velocity) that promptly normalized following ECG-gated optimization of IABP timing parameters. (Fig. 2). Fig. 1Flow of the middle cerebral artery with 1:1 assistance (a) 1:2 assistance (b) and without assistance (c). an increase in the pulsatility index with a drop in diastolic velocity is noted, which improves as ventricular assistance is reduced through the intra-aortic balloon pump. Two representative cardiac cycles are displayed to demonstrate beat-to-beat variability across assistance settings (IABP 1:1, 1:2, paused). Quantitative values in the text were calculated by averaging 3–5 consecutive cycles under the same condition Fig. 2Changes in middle cerebral artery flow velocities after manually adjusting the intra-aortic balloon pump timing, specifically adjusting the deflation time to occur at the beginning of the QRS complex on the ECG, allowing the deflation to continue during early systole. A single representative cycle is shown to emphasize timing landmarks (deflation onset at the QRS) without visual clutter. Quantitative interpretation was based on 3–5-cycle averaging recorded during the same condition METHODS FOR NEUROSONOGRAPHY AND INTERPRETATION: Bedside TCD/TCCD examinations of the MCA were performed via the transtemporal window (typical depth 50–55 mm; sample volume 5–7 mm) with angle correction as needed. Sweep speed was adjusted to visualize multiple cardiac cycles (approximately 3–6 s per screen). For quantitative interpretation, we averaged 3–5 consecutive cycles free of ectopy/artifact at a stable heart rate. Device settings (IABP assistance ratio and timing; VA-ECMO flow and gas) were documented for each acquisition. Figure panels may display 1–2 representative beats solely for visual clarity (e.g., to illustrate beat-to-beat variability under changing assistance, or to mark precise timing landmarks such as QRS-aligned deflation). No numerical values were derived from a single beat; all reported velocities and PI reflect 3–5-beat averaging from the same condition. BEDSIDE INTEGRATION: For patients on VA-ECMO + IABP, we suggest this pragmatic sequence: Interrogate MCA flow with TCD/TCCD, standardizing acquisition (appropriate depth, angle correction as needed, sweep speed showing several cycles). Use 3–5-beat averaging for quantitative interpretation, while selectively using 1–2 displayed beats to highlight beat-to-beat variability or precise timing landmarks [10, 11]. Optimize IABP timing: ensure early diastolic inflation and rapid deflation completed before systole (QRS-gated when needed). Reassess via TCD within minutes to confirm improvement in diastolic velocity and reduction in maladaptive PI trends.Integrate TCD into the multiparametric evaluation for optimizing patient-device interaction. DISCUSSION AND CONCLUSION: Neurological complications during VA-ECMO are multifactorial, stemming from pre-support disease severity, arrest circumstances, ECMO-induced metabolic alterations (especially sudden PaCO₂ changes), residual cardiac function, and supplementary devices [12–14]. Neuromonitoring in these scenarios is crucial for early detection of associated neurological complications and guiding therapeutic interventions. In our case, we observed a temporary decrease in middle cerebral artery (MCA) diastolic flow and elevated pulsatility during 1:1 IABP assistance. This hemodynamic pattern was optimized through careful adjustment of IABP ECG-gated synchronization. The interplay between cardiac function, cerebral perfusion, and intra-aortic balloon pump (IABP) support is mediated chiefly by reductions in effective arterial elastance and left-ventricular afterload with diastolic augmentation; the accompanying rise in cardiac output is typically modest (~ 0.5–1.0 L/min) [5, 6]. At the cerebral level, the balance between diastolic augmentation and systolic impedance shapes the middle cerebral artery (MCA) waveform on TCD. It has been reported—in a small, hypothesis-generating cohort—that during myocardial stunning with very low pulse pressure (PP < 10 mmHg), counterpulsation may reduce cerebral blood flow due to unfavorable timing and limited forward stroke volume; conversely, with partial recovery (PP ≥ 10 mmHg), IABP may increase cerebral flow relative to VA-ECMO alone. Given the small sample size and potential confounders, these observations should not be taken as prescriptive; rather, they reinforce the need for TCD monitoring [15]. This suggests that the cerebral effect of IABP during VA-ECMO is state-dependent and hinges on timing. This is particularly relevant in patients with severe heart failure, where cerebral autoregulation capacity is compromised and cerebral blood flow may be reduced [16]. Ultrasound can be invaluable in the assessment of patients on ECMO, particularly due to the complexity of transporting these patients for diagnostic imaging, allowing for integrated multi-organ monitoring [7]. Transcranial Doppler has been utilized in various clinical scenarios, including the characterization of cerebral perfusion pressure, evaluation of brain death and other neurological assessments, it provides rapid readouts of diastolic velocity and pulsatility trends, captures beat-to-beat responses to timing changes, integrates with NIRS, and can be repeated safely at the bedside where transport is risky [8, 9]. This role is consistent with international recommendations positioning ultrasound-based neuromonitoring as a core ICU competency [10, 11]. The pulsatility index (PI) is not a pure surrogate of distal cerebrovascular resistance; it reflects the interplay among cerebral perfusion pressure, arterial pulse amplitude, vascular compliance and tone, cardiac function and heart rate. Accordingly, PI should be interpreted contextually—with MAP, pulse pressure, and device settings—rather than as a stand-alone estimate of intracranial pressure [17]. In our patient, the maladaptive neurohemodynamic pattern—manifested by rising PI and reduced diastolic velocities—is most consistent with a temporary drop in CPP. Although diastolic reversal was not documented, the observed changes are coherent with a reversal tendency reported in IABP when deflation is delayed [18], and they improved after ECG-gated synchronization. Clinically, the question is seldom “IABP on vs off,” but how the balloon is timed relative to the native cardiac cycle. Inflation should occur at diastole onset to augment aortic diastolic pressure, and deflation should be completed just before systole (often triggered at the QRS) to lower afterload and facilitate forward flow. Suboptimal deflation (too late or slow) may elevate systolic impedance, worsening cerebral perfusion despite adequate mean arterial pressure. In our patient, re-timing deflation restored diastolic augmentation and improved MCA diastolic velocity within minutes. This illustrates how minute-to-minute titration of assistance, guided by TCD, can optimize patient device interaction. LIMITATIONS: This is a single-patient, real-world observation without protocolized hemodynamic targets or invasive intracranial monitoring. TCD measures were interpreted alongside clinical variables, acknowledging that PI reflects multiple determinants. Images were curated for teaching clarity; going forward, we have standardized sweep speed and 3–5-cycle averaging in our protocol to enhance reproducibility.
PMID:39132456 | 0 | 39132456 | PMC11307388 | 10.1016/j.jscai.2024.101337 | fetched | {"title": "The Intra-aortic Balloon Pump: A Focused Review of Physiology, Transport Logistics, Mechanics, and Complications.", "authors": ["Gillespie LE", "Lane BH", "Shaw CR", "Gorder K", "Grisoli A", "Lavallee M", "Gobble O", "Vidosh J", "Deimling D", "Ahmad S", "Hinckley WR", "Brent CM", "Lauria MJ", "Gottula AL"], "journal_abbrev": "J Soc Cardiovasc Angiogr Interv", "year": 2024, "volume": "3", "issue": "5", "pages": "101337", "doi": "10.1016/j.jscai.2024.101337", "pmid": "39132456", "pmcid": "PMC11307388", "pub_types": ["Journal Article", "Review"]} | The Intra-aortic Balloon Pump: A Focused Review of Physiology, Transport Logistics, Mechanics, and Complications. Critical care transport medicine (CCTM) teams are playing an increasing role in the care of patients in cardiogenic shock requiring mechanical circulatory support devices. Hence, it is important that CCTM providers are familiar with the pathophysiology of cardiogenic shock, the role of mechanical circulatory support, and the management of these devices in the transport environment. The intra-aortic balloon pump is a widely used and accessible cardiac support device capable of increasing cardiac output and reducing work on the left ventricle through diastolic augmentation and counterpulsation. This article reviews essential CCTM-based considerations for patients supported by intra-aortic balloon pump, including indications for placement, mechanics and physiology, potential issues during transport, and associated complications. ABSTRACT: Critical care transport medicine (CCTM) teams are playing an increasing role in the care of patients in cardiogenic shock requiring mechanical circulatory support devices. Hence, it is important that CCTM providers are familiar with the pathophysiology of cardiogenic shock, the role of mechanical circulatory support, and the management of these devices in the transport environment. The intra-aortic balloon pump is a widely used and accessible cardiac support device capable of increasing cardiac output and reducing work on the left ventricle through diastolic augmentation and counterpulsation. This article reviews essential CCTM-based considerations for patients supported by intra-aortic balloon pump, including indications for placement, mechanics and physiology, potential issues during transport, and associated complications. INTRODUCTION: Cardiogenic shock (CS) is a state of hypoperfusion secondary to cardiac failure.1,2 CS represents a clinical syndrome, comprising 2 primary categories: acute myocardial infarction (AMI)-CS or acute decompensated heart failure (ADHF)-CS. These 2 subcategories carry distinct clinical and hemodynamic characteristics.3,4 Prevalence of CS is increasing, with 8% patients of ST-elevation myocardial infarction developing CS.5 The rising prevalence of CS has resulted in increasing utilization of mechanical circulatory support (MCS). Although other mechanical circulatory support devices have been developed, the intra-aortic balloon pump (IABP) remains the most commonly used.6 It has specific advantages in its ease of insertion and is an attractive option in hospitals with limited resources. Patients supported by an IABP may require transport by critical care transport medicine (CCTM) teams to facilities with more advanced intervention or further care options. The successful interfacility transport of patients supported by IABP has been described, although in limited detail and scope.7, 8, 9 To our knowledge, there are no professional society guidelines specifically describing the transport of patients with an IABP. This review focuses on the essential CCTM-based considerations for patients supported by IABP, including indications for placement, mechanics and physiology, and associated complications. DEFINITION OF CS: Major clinical trials have used variable criteria to identify patients with CS, leaving clinicians without a standardized definition (Figure 1). However, most recently, the Shock Academic Research Consortium group, comprising multidisciplinary expert panels, suggested both a practical and a clinical trial definition of CS.10 From a clinical practice perspective, their suggested definition for CS is as follows: “Cardiac disorder that results in both clinical and biochemical evidence of sustained tissue hypoperfusion,” whereas the research-focused definition suggests “Cardiac disorder that results in a systolic blood pressure (SBP) <90 mm Hg for ≥30 minutes (or the need for vasopressors, inotropes, or MCS to maintain systolic blood pressure of ≥90 mm Hg) with evidence of hypoperfusion.”10Figure 1SCAI classification of cardiogenic shock incorporating the 2022 revisions.11,12 BNP, brain natriuretic peptide; CPR, cardiopulmonary resuscitation; CVP, central venous pressure; JVP, jugular venous pressure; LFT, liver function test; MAP, mean arterial pressure; PCWP, pulmonary capillary wedge pressure; SBP, systolic blood pressure. Similarly, a multidisciplinary group was convened by the Society for Cardiovascular Angiography & Interventions (SCAI) to establish a novel CS classification system. The SCAI criteria is a letter-based (A through E) system using a combination of physical examination, clinical factors, biochemical markers, and hemodynamic parameters to classify patients into separate categories (Figure 1).11 There are many advantages to the SCAI criteria, including applicability to previous trials and retrospective data, facilitating further concerted research efforts.11 These criteria allow clinicians to make rapid, bedside assessments as patients migrate between categories of severity rather than relying on more resource-intensive data. An update to the original SCAI criteria was published in 2022, with modifications aimed at clarifying differentiating factors separating categories, and rephrasing the physical examination, and biochemical and hemodynamic parameters to include subcategories.11,12 Multiple validation studies have used the SCAI criteria for CS, which support the correlation between mortality and increasing SCAI grade.12 Critical care transport teams may use SCAI grade in communication with referring or receiving facility providers and assessment of patient deterioration and potential interventions. Overall, the definition of CS remains heterogeneous, whether through landmark clinical trials or in bedside and clinical practice.10,13, 14, 15 PRESSURE-VOLUME LOOPS AND CS: The pressure-volume (PV) loop (Figure 2) serves as a model demonstrating both normal cardiac physiology and the effects of maladaptive mechanisms in CS. There are 4 distinct stages of the cardiac cycle demonstrated on the PV loop: (1) ventricular filling, (2) isovolumetric contraction, (3) ejection, and (4) isovolumetric relaxation.Figure 2Pressure-volume loop in normal cardiac physiology. The PV loop demonstrates normal cardiac physiology and the relationships between volume and pressure. Diastolic filling of the ventricle (D to A) is followed by early systole and isovolumetric contraction (A to B), during which the pressure generated by the ventricle results in opening of the semilunar valves, closure of the atrioventricular valves, and forward ejection of blood. Systolic ejection (B to C) is followed by ventricular isovolumetric relaxation (C to D) during which ventricular pressures drop below the diastolic aortic and pulmonary pressure leading to semilunar valve closure and the beginning of diastole. EDPVR, end-diastolic pressure-volume relationship; ESPVR, end-systolic pressure-volume relationship; LV, left ventricle. Correlating to distinct phase changes in the cardiac cycle, at point A, end-diastolic filling has occurred. Extrapolation of this portion of the cycle tracing demonstrates the end-diastolic PV relationship. This nonlinear relationship describes innate structural properties of the left ventricle (LV), which are affected by pathologies such as hypertrophy, myocyte remodeling, fibrosis, and ischemia.16 For example, the slope change (diastolic pressure/diastolic volume) reflects stiffness of the LV, and the inverse of this equation reflects compliance (diastolic volume/diastolic pressure). The vertical line from A to B represents isovolumetric contraction, with subsequent aortic valve (AV) opening secondary to the gradient between the intraventricular pressure and aortic diastolic pressure. From point B to point C, blood is ejected out of the LV into the systemic circulation. When linearly extrapolated relative to the volume axis intersection (V0), point C in the cardiac cycle represents end-systolic pressure-volume relationship (ESPVR). ESPVR is the maximum pressure generated for a given ventricular volume, corresponding to contractility. The relationship between the slope of this line (end-systolic elastance Ees) and the volume axis intersection (V0) is reflected as follows: end-systolic pressure (Pes) = Ees × (Ves − V0), and demonstrates ESPVR.16 Ea depicts the afterload. The Ea line starts on the volume axis at the end-diastolic volume and intersects the ESPVR at the ventricular end-systolic PV point of the PV loop. As contractility increases, the slope of this line is translated leftward and/or upward (Figure 2). At point C, isovolumetric relaxation occurs, which permits mitral valve opening (point D) and the subsequent initiation of the upcoming cardiac cycle. THE MALADAPTIVE, SELF-PERPETUATING CYCLE OF CS: Cardiogenic shock generally results from impaired myocardial contractility, leading to a maladaptive spiral of reduced cardiac output (CO), systemic hypoperfusion, and coronary ischemia. Propagation of coronary ischemia further reduces contractility. This self-perpetuating cycle activates multiple neurohumoral pathways resulting in a vasoconstrictive and volume overload state, which ultimately progresses to multiorgan failure and death.17,18 The PV loop can aid in describing the hemodynamic changes in the setting of CS. The effects of acute LV dysfunction are shown in Figure 3. The changes in the PV loop include downward and rightward shift of ESPVR reflecting reduced contractility, with a mild increase in end-diastolic PV relationship as filling pressures are elevated. The PV loop area, corresponding to stroke work (ventricular work required for ejection of stroke volume), is reduced.Figure 3Neurohumoral activation and LV dysfunction impact the PV loop. The physiologic adaptations found in cardiogenic shock with acute left ventricular dysfunction are demonstrated by changes in the PV loop. With a downward and rightward shift in the acute LV dysfunction loop, filling pressures (LV volume) are significantly elevated, with lower overall generated systolic pressures, and reduced contractility as portrayed by the shift in ESPVR. Neurohumoral activation releases endogenous catecholamines resulting in increased myocardial oxygen demands, centralization of blood volume leading to pulmonary congestion, and further elevation of filling pressures. LV, left ventricle; PV, pressure-volume. As demonstrated in the PV loop, increased left-sided filling pressures result in pulmonary congestion, hypoxic vasoconstriction, increased right ventricular afterload, and elevated central venous pressure. In response to this hypoperfusion, endogenous catecholamines are released as a compensatory mechanism. This neurohumoral activation results in an increase in contractility and heart rate, increasing myocardial oxygen demand, further taxing an overstretched cardiac system.2,19 Finally, to support a perfusing arterial pressure, venous and arteriolar vascular beds constrict, leading to centralization of blood volume, further elevating filling pressures.19 IABP IMPACT ON CARDIAC FUNCTION: The well-described maladaptive physiologic mechanisms implicated in CS serve as potential opportunities for targeted interventions, such as MCS devices. These devices, when applied to the right patient, have the ability to counteract some of these maladaptive processes by offloading the LV, reducing afterload, and augmenting coronary perfusion pressure (CPP). However, the response to MCS device insertion varies by CS subpopulation, whether AMI-CS or ADHF-CS given the distinct compensatory pathophysiology and hemodynamics.20, 21, 22 There is significant active research interest in studying the impact of the IABP and MCS on patients with ADHF-CS, with some studies suggesting benefit in this subpopulation. Given the response variability in individual patients, there may be responders and nonresponders to IABP therapy.20, 21, 22 The IABP impacts CO and coronary perfusion. Immediately before systole (QRS-T interval), the balloon is rapidly deflated. This rapid deflation results in a decrease in afterload and increase in LV stroke volume, in addition to decreased peak systemic blood pressure, earlier opening of the AV, and a shorter LV isometric contraction phase (Figure 4).23 In diastole (T-P interval), the balloon inflates, temporarily increasing afterload and ultimately increasing the CPP and systemic diastolic blood pressure.23 CPP is the difference between aortic end-diastolic pressure and left ventricular end-diastolic pressure (LVEDP). Therefore, the increase in CPP is mediated by an increase in aortic pressure with a concomitant decrease in LVEDP.23 However, it is vital to note that this increase in CPP does not translate necessarily into increased coronary blood flow (CBF) because CBF is also impacted by coronary vascular resistance.23,24 In addition to increasing CPP, the IABP provides a modest increase in CO of 0.5-1 L/min, increases myocardial oxygen delivery, and decreases myocardial oxygen demand.23, 24, 25, 26Figure 4IABP impact on the PV loop. The IABP has multiple favorable effects on cardiac function. While inflated, the IAB displaces intra-aortic blood during diastole, propagating forward movement of blood into systemic circulation and ideally, increasing coronary blood flow. Owing to the reduction of end-diastolic intra-aortic blood volume, the pressure required of the LV to initiate aortic valve opening is reduced (afterload reduction, seen at point B), leading to reduced LV work and myocardial oxygen demand. Before systole, the balloon then deflates, which also contributes to reduced LV afterload and aortic pressures due to negative pressure mechanics. The time spent in isovolumetric contraction is reduced, leading to reduced myocardial demand. SV increases (EDV, point A; ESV, point C), resulting in modest increases in CO. EDV, end-diastolic volume; IAB, intra-aortic balloon; IABP, intra-aortic balloon pump; LV, left ventricle. IABP MECHANICS: Studies first demonstrated the utility of the IABP and counterpulsation in 1958, and the device was used clinically just 1 decade later. The IABP facilitates diastolic augmentation to increase CBF and LV systolic offloading to decrease LV work.27 The IABP consists of a console, which is responsible for delivery of a volume of gas, catheter, and balloon (Figure 5).23, 24, 25Figure 5IABP console configuration. The IABP console consists of a foldable screen, where settings may be entered and adjusted, waveforms are represented, and alarms display; a connection panel, for fiberoptic (top right port), ECG (green, top left), pressure/arterial line (second from top left, red), and helium extender tubing (donut port, central) cables; 2 lithium ion batteries (pictured below front-facing vents); and a helium tank system (not pictured, located on the lower back-side of the console). ECG, electrocardiogram; IABP, intra-aortic balloon pump. The intra-aortic balloon has a uniform construction, although there are variations in membrane length (22.0-27.5 cm), inflated diameter (15-18 mm), and balloon volume (30-50 mL).24 The balloon is rapidly inflated and deflated by injection and retrieval of a gas, typically helium or carbon dioxide.24 Helium, which is inert, is the preferred agent, because its lower viscosity reduces resistance and increases laminar flow, allowing for more rapid gas entry and retrieval.24 Optimal balloon sizing requires that no more than 90% of the aortic diameter is obstructed when the balloon inflates, although in clinical practice, the aortic diameter is infrequently measured before insertion.24 Improper sizing carries risks as oversized balloons increase the risk of vascular complications, whereas an undersized balloon will offer suboptimal cardiac augmentation.24 In the majority of cases, the 40-mL balloon is appropriate and should be used in patients 162-182 cm tall.24,25 Use of a 50-mL balloon should be considered in patients who are >182 cm.25 IABP CONSOLE: The console serves as the control unit for the device and is responsible for inflating and deflating the balloon in a coordinated manner. It consists of a monitor unit, control unit, gas source, and valve unit (Figure 5).23, 24, 25 The monitoring unit receives the data from the patient (ie, electrocardiogram [ECG] and infrared blood pressure monitor), and the control unit uses this information to inject and retrieve the gas.23,24 The monitoring unit displays the systemic and augmented blood pressure and allows for monitoring of the device’s function. The control unit allows for modulation of the balloon’s ratio of supported beats. The number of supported beats is inversely proportional to the support that the balloon is providing, that is, if the balloon is in a 1:1 mode, each beat is being supported, whereas in 1:2 mode, only every second beat is supported.28 The monitoring unit also permits observation of the device’s functionality. A normal tracing of a 1:1 IABP mode can be seen in Figures 6 and 7.28 The user will be able to observe an incorrectly triggering IABP by looking at the device’s monitor tracing or the waveform of the patient’s arterial line. The control unit will read the data from the monitoring unit and inject gas into the balloon during the T-P interval of the ECG and will deflate during the QRS-T interval when the ECG trigger mode is used.14,23, 24, 25, 26,29Figure 6Pictorial representation of IABP in systole and diastole with ECG tracing, arterial augmented waveforms and corresponding IABP tracings. The IABP functions to inflate during diastole and deflate during systole. This can be accomplished by timing relative to ECG or pressure waveforms, pictured above, to accurately inflate during the appropriate portion of the cardiac cycle. The IAB waveform, pictured in blue, is timed to correlate vertically with diastole in the arterial and ECG tracings. ECG, electrocardiogram; IAB, intra-aortic balloon; IABP, intra-aortic balloon pump.Figure 7Normal and abnormal IABP waveforms and their impact on cardiac physiology. IABP monitor interface demonstrating waveform changes in the setting of (A) correct inflation timing, (B) early inflation, (C) early deflation, (D) late inflation, and (E) late deflation. The arterial waveform in normal IABP inflation timing results in augmentation of diastole correlating with balloon inflation. The balloon pressure waveform illustrates rapid inflation of the balloon, followed by the plateau at maximal inflation, and subsequently rapid deflation before returning to baseline in preparation for the next cycle. In early inflation (B), afterload increases, and the aortic valve undergoes premature closure, which increases the end-systolic volume and reduces CO. In early deflation (C), there is reduced overall impact and efficacy of the IABP on cardiac function; there is minimal diastolic augmentation and a lack of reduction in myocardial oxygen demand and LV afterload. In late inflation (D), there is significantly reduced diastolic augmentation given the delayed timing; there still may be a small impact on improved coronary perfusion, although it is suboptimal compared with the potential of an adequately timed device. Finally, in late deflation (E), there is an increase, rather than decrease, in LV afterload, which translates into increased myocardial oxygen demand. CO, cardiac output; IABP, intra-aortic balloon pressure. IABP INSERTION TECHNIQUE: The IABP is most commonly implanted through the femoral artery; however, alternative sites include the brachial, axillary, subclavian, or iliac arteries.23, 24, 25,28,30 Longer-term IABP implants may also be placed in the operating room, using surgical cutdown. The IABP is typically placed under fluoroscopic guidance but can be inserted at bedside as well. Proper placement should be confirmed by chest radiograph before use. A radiopaque metallic marker on the distal tip of the IABP is used to assess placement and should be ∼2 cm superior to the carina, which approximates the area of the takeoff of the left subclavian artery in most adults. An alternative landmark that in clinical practice is frequently used is the aortic knob, with the IABP tip located 1-2 cm distal to the aortic knob.23, 24, 25,28 Of note, some operators may place the IABP using a 9F sidearm sheath, which allows for pressure transduction of the sidearm, serving as a backup in the event of failure of the primary pressure measuring system. Once the balloon’s location has been confirmed, the gas lumen of the balloon must be checked to ensure the absence of blood, which would suggest balloon rupture. Once completed, the console is set and therapy initiated.25 EVIDENCE FOR IABP UTILITY IN THE LITERATURE: Despite being the most commonly used device for CS, the IABP lacks reproducible evidence supporting improvement in patient-centered outcomes, specifically reduction in mortality.14,23,24,29 However, the use of the IABP has been shown to impact physiologic parameters in a statistically significant manner: an increase in cardiac index, decrease in pulmonary capillary wedge pressure, and decrease in multiorgan dysfunction have been observed with its use.29 Initial studies between the 1970s and 1990s reported a mortality benefit in patients with CS; however, subsequent trials have shown consistently negative results.23 IABP-SHOCK, and the follow-up study IABP-SHOCK II, found no differences in mortality, hemodynamics, or vasopressor requirements in patients with AMI and/or CS who were medically managed with or without augmentation via an IABP.23,24,29 Lack of benefit in this patient population has been substantiated by numerous trials and meta-analyses.23,29 However, there are data that suggest potential utility of the IABP in patient populations without AMI-CS, specifically, in patients affected by ADHF.22,31, 32, 33 Patients in CS secondary to ADHF are characterized by distinct hemodynamic profiles in comparison with patients with AMI-CS. For example, 1 retrospective study compared hemodynamic responses to IABP in patients with and without AMI-CS. Before IABP insertion, the ADHF-CS group had higher baseline pulmonary artery pressures than the AMI-CS group, and after IABP insertion, they showed a nearly 5-fold greater augmentation in CO, whereas patients with AMI-CS experienced almost no improvement in CO.20,32 The IABP may also serve as a bridge to advanced heart failure therapies in patients with advanced heart failure.22,31,33 There is a need for larger randomized controlled trials evaluating the potential benefit of IABP in patients with ADHF-CS. Irrespective of the primary indication, in many locales, the IABP remains the most frequently inserted MCS device, and CCTM crews will frequently transfer these patients. ECG TRIGGER: The IABP is meant to inflate during the closing of the AV and deflate during its opening. The ECG trigger option is the preferred and generally the most reliable. In this mode, the R wave on the patient’s ECG corresponds with the opening of the AV and onset of systole; this triggers deflation of the IABP. The IABP then senses the middle of the T wave on the patient’s ECG, corresponding with diastole, and triggers inflation. Arrhythmias can impair the IABP’s ability to trigger inflation, as can low ECG voltage, improper electrode placement, and electrical interference with the ECG signal. It is important to ensure proper electrode placement and that the lead with the most prominent R wave is selected and fixing underlying causes of arrhythmia (ie, administer appropriate anti-arrhythmics, defibrillate, and correct electrolytes). If the issue persists, the IABP can be set to a different trigger option. ATRIAL FIBRILLATION TRIGGER: If the patient is in atrial fibrillation, an atrial fibrillation trigger mode can often be used, which triggers balloon deflation at each R wave, despite variability between the complexes. PRESSURE TRIGGER: Often used when ECG quality is poor, the arterial waveform can also trigger inflation and deflation of the IABP. In this mode, the dicrotic notch in the arterial waveform triggers balloon inflation, and the point just before systolic upstroke (corresponding to the opening of the AV) will trigger deflation. This mode can also be used during cardiopulmonary resuscitation because the arterial waveform should be generated by effective chest compressions. INTERNAL TRIGGER: The IABP can be set to an intrinsic/asynchronous rate without physiologic triggering. This mode should only be used in situations where CO and ECG tracings are absent, as in the case of asystole or in a patient supported on cardiopulmonary bypass. PACER TRIGGER: Pacing spikes (if present) can also be used to trigger IABP inflation. The IABP is triggered to inflate by the V pacing spikes. This mode can be safely used if the patient is 100% pacer dependent. AUTOMODE: Automode is a feature that is capable of automatically discerning and choosing the most appropriate ECG lead and trigger and sets inflation and deflation timing. This mode also is able to respond and select the appropriate trigger in arrhythmias (eg, atrial fibrillation). The simplicity and safety of automode makes this generally the best operation mode choice for transport providers, minimizing potential for error. In comparison, semi-automode requires that the provider establishes timing, then the device console automatically adjusts subsequent timing to the patient’s heart rate and rhythm, and that the provider selects the lead and trigger sources. IABP FAILURE ALARM: This alarm suggests an intrinsic failure within the IABP console or computer. Consider reinitiating the console or computer device or changing devices if necessary after troubleshooting. IABP DISCONNECTED ALARM: This alarm suggests disconnection of the IABP extension tubing. Ensure IABP extension tubing is connected and refilled with gas. RAPID GAS LOSS ALARM: This alarm suggests a leak within the IABP circuit or a hole or leak in the balloon. Because the IABP is a closed system, a leak may cause the balloon’s waveform to begin to fall below the baseline of zero. Some manufacturers may have a “pump rupture” or “autofill failure” alarm instead. Verify connections are without leaks. Examine the catheter for blood or coffee ground-like material. If this is observed, the integrity of the balloon has been compromised and likely needs to be removed. Stop the pump and prepare for removal and replacement on arrival to the receiving center. CCTM providers should be prepared to increase support for patients who lose the function of their IABP en route. CHECK IABP CATHETER ALARM: This alarm suggests a kink in the IABP tubing or catheter. If a sheathed IABP was placed, it is possible that the IABP is partially retained in the sheath, hampering inflation of the balloon. This alarm could also sound if the IABP has not fully unfolded in the aorta. The balloon pressure waveform may have a rounded peak rather than a flat plateau. Check the IABP tubing and catheter, removing any kinks. If a sheathed IABP was placed, check the markings on the IABP to determine whether displacement has occurred. Transport providers should not manipulate the IABP insertion depth and, if displaced, should notify the receiving facility providers. LOW HELIUM ALARM: This alarm suggests the helium supply is running low. Replace the helium cylinder. Ensure there are backup cylinders available to avoid prolonged pauses of the IABP between cylinder changes. Some manufacturers may have an “autofill failure” alarm instead. Certain IABP models provide for rapid helium refilling when the pump console is separated from the cart, as in most transport scenarios; helium may be rapidly refilled within minutes if the console is reconnected to a same-model cart. AUGMENTATION BELOW LIMIT SET ALARM: This alarm suggests either augmentation is set too high or there has been a change in the patient’s hemodynamic status. Assess the patient for decreasing cardiac function or worsening vascular resistance and add inotropes or vasoactive agents as needed. Assess for improving cardiac function and consider decreasing or removing IABP if appropriate. Consider decreasing the alarm limit if other clinical data suggest the patient is adequately supported. NO TRIGGER ALARM: This alarm suggests that the IABP is unable to sense the trigger mode that has been selected. There are various trigger modes that may be useful in different circumstances, detailed in the earlier section. Troubleshooting will be based on the triggering mode being used and may require changing to a different mode. TIMING ERRORS: Correct timing of the inflation and deflation cycles of the IABP is critical. If the balloon is not adequately timed to inflate and deflate in sync with the patient’s cardiac cycle, it can impair an already dysfunctional LV. Timing is assessed in the 1:2 mode so that pressure differences between augmented and nonaugmented beats can be observed and used for diagnostics. A correctly timed IABP should result in an arterial line waveform that shows inflation occurring at the dicrotic notch and forming a sharp V shape. The slope of the augmented diastolic pressure should be parallel to, and reach a peak higher than, the previous systolic upstroke. The assisted end-diastolic blood pressure (EDBP) should be lower than the previous unassisted EDBP by 15-20 mm Hg. Finally, the peak of assisted SBP that follows IABP inflation should be ∼5 mm Hg lower than the previous unassisted SBP. Errors with timing, causing either early or late inflation or deflation, can be detected by variance on the arterial waveforms. EARLY INFLATION: Early inflation will cause the IABP to inflate before the AV has closed. This has the potential to cause the AV to close prematurely, because of increased afterload, leading to impaired LV emptying and possibly aortic regurgitation. Inc
PMID:39132456 | 1 | 39132456 | PMC11307388 | 10.1016/j.jscai.2024.101337 | fetched |  | omplete LV emptying will lead to an increase in LV end-diastolic volume, LVEDP, and pulmonary capillary wedge pressure, which will increase myocardial oxygen demand. Early inflation will lead to diastolic augmentation occurring prior to the dicrotic notch with loss of the desired “V” shape. The diastolic augmentation may encroach onto the unassisted systolic peak to such a degree that the 2 peaks become difficult to differentiate. Early inflation can be corrected by adjusting the timing so that the IABP inflates at the dicrotic notch; however, the best solution is to use automode to optimize timing rather than manually manipulating this variable in critically ill patients. LATE INFLATION: Late inflation of the IABP occurs well after closure of the AV. This has the potential to cause inadequate perfusion of the coronary arteries due to suboptimal diastolic augmentation. The arterial waveform will show IABP inflation occurring well after the dicrotic notch, again resulting in loss of the optimal V shape. The augmented diastolic blood pressure may be equal to or lower than the unassisted SBP. This can be corrected by adjusting the timing so the IABP inflates at the dicrotic notch; however, similar to the suggestion for remedying early inflation, automode should be used rather than transport providers manually altering variables on the console. EARLY DEFLATION: Early deflation of IABP during the diastolic phase ends before isovolumetric contraction occurs. This decreases the efficacy of the IABP to reduce afterload, reducing the boost in CPP and obviating the potential reduction in myocardial oxygen demand. Early deflation may lead to retrograde coronary and carotid blood flow due to the vacuum-like effect in mid-diastole. The arterial waveform will show a sharp drop in pressure soon after the peak of diastolic augmentation. The assisted SBP may be equal to or higher than the unassisted SBP and the assisted EDBP may be equal to or less than the unassisted EDBP. Early deflation can be corrected by prolonging the IABP inflation time, although automode typically accounts for and corrects this error. LATE DEFLATION: Deflation should end near the end of diastole, during isovolumetric contraction; if IABP deflation is prolonged, it may still be in the process of deflating after the AV has opened. In this scenario, rather than reducing afterload, the resistance from a still-inflated balloon during systole will increase afterload, interfering with LV ejection, and prolonging isovolumetric contraction. These effects will increase myocardial oxygen demand. The characteristic arterial waveform of late deflation will show an assisted end-diastolic pressure that is equal to or greater than the unassisted end-diastolic pressure (rather than appropriately decreased) and a widened augmented diastolic peak. If late deflation occurs, the IABP inflation time should be shortened, although automode typically accounts for and corrects this error. There are several alarms common to various IABP brands. It is important for the provider to be able to recognize the potential causes of these alarms and to address them. CARDIAC ARREST: In the setting of cardiac arrest, IABP support may be continued. The ECG trigger becomes dysfunctional and cannot be used as the primary timing source; in shockable rhythms, there is no discernible organized rhythm. Interestingly, in situations where the patient is in pulseless electrical activity, the IABP will continue to sense the organized rhythm; however, during compressions, there will be significant artifact. The pressure trigger, rather than ECG trigger, is the preferred sensing method in this scenario. In the transport setting with a small number of providers, this can be accomplished by simply disconnecting the ECG leads from the pump console. The patient’s underlying cardiac rhythm should still be monitored through the primary cardiac monitor using electrodes or pads. Defibrillation can be safely performed with the IABP device functioning. Tachydysrhythmias may also impact IABP performance. Small prospective and some retrospective studies have evaluated the impact of tachycardia on IABP efficacy and function, suggesting reduced performance indices at heart rates greater than 110 bpm; however, there are some IABP models that are capable of matching 1:1 rates up to 220 bpm.34, 35, 36 An opportunity for troubleshooting with reduced IABP efficacy is to change the ratio from 1:1 to 1:2. The primary treatment for tachydysrhythmias is to treat the underlying cause. TRANSPORT CONSIDERATIONS: When CCTM teams receive a transport request from a cardiac catheterization laboratory or other critical care area, we recommend that the transport team routinely inquire about the use of MCS devices. Patients supported by IABP will require additional preparation to ensure safe transport, including additional specialized equipment, appropriately trained personnel, and any necessary adaptations for the transport method (eg, a floor securement device for rotor-wing transport). EQUIPMENT ASSESSMENT: To ensure adequate equipment operations and safety, the following should be performed:1.Ensure the pump battery is fully charged.a.Current commercial devices support battery life of 2 to 2.25 hours.37 A backup battery should be available in case of equipment failure or unforeseen transport delay.b.The pump should be connected to AC power or an inverter whenever possible to preserve battery life.c.The mechanic should ensure the vehicle’s electrical system can accommodate pump power requirements along with other necessary electric systems so the system will not be overloaded in transport.2.Check the helium tank level.a.Before departure, the helium tank should be filled. For air transport, the balloon will empty and fill with some variation in conjunction with altitude changes.3.Secure the pump console onto the transport vehicle mount to comply with local and national safety standards.a.Commercial devices include portable consoles to meet the physical constraints of medical transport. These should be installed according to manufacturer recommendations and comply with ground and aviation safety standards.b.All personnel should be familiar with proper device installation within the transport vehicle.c.During patient transport, the controller screen should be visible to the medical team to monitor function and view alarms, which may not be audible with ambient background noise.4.Include adapters for connection to devices from local facilities.a.The transport service should communicate with local cardiac catheterization facilities regarding devices and compatibility.b.The most common commercial devices such as Arrow IABP (Teleflex) and Maquet/Datascope IABP (Getinge) include adapter extension catheters to connect the balloon to different consoles.c.Bring any catheter extension tubing from the transferring facility to the receiving facility.5.Be prepared for infrequently encountered but severe complications with backup equipment.a.Include a 60-mL syringe with a 3-way stopcock for manual rapid balloon inflation and deflation of 40-50 mL of air through the helium extender tubing, in case of complete console dysfunction and failure. The goal here is no longer hemodynamic support but rather the prevention of thrombosis. PATIENT ASSESSMENT: In addition to the standard history and examination before transport of any critically ill patient, assessment of a patient with an intra-aortic balloon pump should include evaluation of the device, functioning of the device, and consideration of potential complications. Placement of the device is radiographically confirmed after the procedure. Standard intra-aortic balloons have a radiopaque tip that should appear ∼1-2 centimeters distal to the aortic knob (Figure 8). If there are any concerns about placement or migration of the device on assessment, a repeat x-ray and/or evaluation by the referring physician should be requested by the CCTM team before transport.Figure 8Chest x-ray demonstrating correct positioning of the IAB. The IAB is highlighted by the radiopaque marker at the distal point of the balloon (superior aspect of the yellow-highlighted portion). There is also a radiopaque midline marker within the balloon itself, which is faintly visible. Additional support devices pictured include bilateral chest tubes, an endotracheal tube, and a right internal jugular approach Swan-Ganz catheter. IAB, intra-aortic balloon. The most common complications of IABPs include limb ischemia and hemorrhage.38 Therefore, a thorough neurovascular exam must be conducted on all patients before transport and again at the receiving facility. Distal pulses should be assessed, with Doppler if necessary, to monitor for limb ischemia. Loss of the left radial pulse is concerning for limb ischemia due to either thrombosis or proximal migration of balloon to occlude the left subclavian artery. The insertion site should be examined for ecchymosis, hematoma, or bleeding. If feasible, leaving the insertion site visible and easily accessible will facilitate rapid site assessment and IABP repositioning throughout transport if needed. When transferring a patient with an IABP in place, do not raise the head of the bed greater than 30° because this may strain or kink the femoral catheter and cause bleeding. To maintain and maximize static positioning of the device, there are several practical techniques that may be considered, such as use of foley anchors for device securement, or knee immobilizers to reduce the potential for patient movement kinking or otherwise impacting device and line positioning. Before transferring the patient’s device to the transport pump console, note settings, and timing. If the patient’s device is set at an augmentation ratio of 1:1, adjust to 1:2 to analyze timing and augmentation, and then return to original settings. A standard checklist for patient assessment and documentation serves to ensure essential measures, and evaluations are recorded and performed. Examples of patient assessment checklists and IABP device evaluation forms for transport are provided in Supplemental Figures S1 and S2, respectively. ALTITUDE CHANGES: For proper counterpulsation therapy and to prevent overdistention, the IABP must be adjusted to compensate for changes in atmospheric pressure. This is an essential consideration in the context of aeromedical transport of patients in rotary-wing or fixed-wing aircraft. Some IABPs will perform this compensation automatically. For example, in the autofill mode, the Maquet CS300 system (Getinge) will automatically purge and fill the balloon when the atmospheric pressure decreases by 25 mm Hg (occurring approximately every 1000 feet [304.8 m] of ascent) or when the atmospheric pressure increases by 50 mm Hg (occurring approximately every 2000 feet [609.6 m] of descent). If the device fails to automatically adjust, pressing “pump status on” will cause the pump to automatically adjust to altitude. Rapid ascent without appropriate adjustment will cause “high baseline” or “high fill pressure” alarms, terminating counterpulsation. Rapid descent may cause “helium loss” or “gas loss” alarms.39 CCTM providers should be aware of the compensatory capabilities of the IABP system used by their agency. LOGISTICAL CONSIDERATIONS: Patients supported by MCS, such as an IABP, challenge the spatial constraints of CCTM, with the most restrictive conditions being the helicopter environment. However, lighter and more compact modern equipment have made interfacility air transport of patients with IABP relatively common. Figure 9 depicts a potential cabin configuration of an EC145 helicopter with a patient supported by an IABP in which the flight physician or paramedic sits at the patient’s head while a flight nurse sits to the patient’s left side with the IABP console. Any configuration should ensure accessibility to infusion pumps, the IABP console, cables/connections into the IABP, patient monitors, and a ventilator, if the patient is supported by either invasive or noninvasive positive pressure ventilation, because patients experiencing CS often require in-flight adjustments of both medical and mechanical support.Figure 9Top-down view of the configuration of an EC145 helicopter with a patient supported by an IABP. Pictured is the proposed, commonly encountered configuration for rotor-wing transport of a patient with IABP support. Dotted lines represent closed doors, with clamshell doors depicted in the open position. CONCLUSION: CS is characterized, in general, by impaired ventricular contractility leading to hypoperfusion, which then promotes further myocardial insult and dysfunction, ultimately culminating in systemic end organ damage. MCS devices, including the IABP, can be used to supplement perfusion and avoid irreversible end organ damage in this population. The IABP is a percutaneously inserted counterpulsation pump positioned in the descending aorta that serves to improve coronary perfusion, reduce afterload, and augment CO by approximately 0.5-1 L/min. As IABPs are increasingly encountered in the transport environment, knowledge of the underlying physiologic changes, mechanics, and troubleshooting procedures for the IABP are essential (Central Illustration). The initial assessment should include a diligent physical examination, which serves to ensure device securement, assessment for device complications, and inspection of device settings. Careful monitoring for complications, such as inflation and deflation timing errors, balloon rupture, and helium gas supply errors is an essential component of care and requires specific training.Central IllustrationThe intra-aortic balloon pump is a frequently encountered mechanical support device. Critical care providers must be expertly trained in caring for patients supported by the IABP, given the essential role of transport providers in connecting patients to highly resourced specialty centers for cardiogenic shock. IABP, intra-aortic balloon pressure. The transport of patients supported by IABP devices has been successfully and safely accomplished with well-trained providers and systems. The provision of high-quality critical care transport in patients with IABP devices requires diligence and in-depth comprehension of the native and device-impacted physiology and intratransport complications.
PMID:36263076 | 0 | 36263076 | PMC9536734 | 10.36628/ijhf.2019.0015 | fetched | {"title": "Mechanical Circulatory Support for Acute Heart Failure Complicated by Cardiogenic Shock.", "authors": ["Choi MS", "Shim H", "Cho YH"], "journal_abbrev": "Int J Heart Fail", "year": 2020, "volume": "2", "issue": "1", "pages": "23-44", "doi": "10.36628/ijhf.2019.0015", "pmid": "36263076", "pmcid": "PMC9536734", "pub_types": ["Journal Article", "Review"]} | Mechanical Circulatory Support for Acute Heart Failure Complicated by Cardiogenic Shock. Acute heart failure is a potentially life-threatening condition that can lead to cardiogenic shock, which is associated with hypotension and organ failure. Although there have been many studies on the treatment for cardiogenic shock, early mortality remains high at 40-50%. No new medicines for cardiogenic shock have been developed. Recently, there has been a gradual decline in the use of the intra-aortic balloon pump mainly due to a lack of adequate hemodynamic support. Extracorporeal membrane oxygenation and the percutaneous ventricular assist device have become more widely used in recent years. A thorough understanding of the mechanisms of such mechanical support devices and their hemodynamic effects, components of the devices, implantation technique, management, criteria for indications or contraindications of use, and clinical outcomes as well as multidisciplinary decision making may improve the outcomes in patients experiencing cardiogenic shock. ABSTRACT: Acute heart failure is a potentially life-threatening condition that can lead to cardiogenic shock, which is associated with hypotension and organ failure. Although there have been many studies on the treatment for cardiogenic shock, early mortality remains high at 40–50%. No new medicines for cardiogenic shock have been developed. Recently, there has been a gradual decline in the use of the intra-aortic balloon pump mainly due to a lack of adequate hemodynamic support. Extracorporeal membrane oxygenation and the percutaneous ventricular assist device have become more widely used in recent years. A thorough understanding of the mechanisms of such mechanical support devices and their hemodynamic effects, components of the devices, implantation technique, management, criteria for indications or contraindications of use, and clinical outcomes as well as multidisciplinary decision making may improve the outcomes in patients experiencing cardiogenic shock. ACUTE HEART FAILURE AND CARDIOGENIC SHOCK: Acute heart failure (AHF) and cardiogenic shock (CS) are similar but somewhat different concepts. The 2016 European Society of Cardiology (ESC) described AHF as a “rapid onset or worsening of symptoms and/or signs of heart failure”.1) CS is defined as a state of circulatory shock caused by dysfunction of cardiac pumping. We believe that the cause of both conditions is dysfunctional cardiac pumping. However, the presentation is more severe in CS than in AHF. Circulatory shock refers to hypotension accompanied by symptoms and signs of organ hypoperfusion such as altered mentality, cold and mottled skin, reduced urine output, low central venous oxygen saturation, and increased serum lactate.2) Symptoms and signs of heart failure are dyspnea, indigestion, peripheral edema, and increased jugular venous pressure and so on. The causes of AHF and CS are diverse. Acute myocardial infarction (AMI) with subsequent ventricular dysfunction, the most frequent cause of CS, accounts for approximately 80% of cases. Mechanical complications of AMI, such as ventricular septal (4%) or free wall (2%) rupture and acute severe mitral regurgitation (7%) are less frequent causes of CS after AMI.3) Causes of non-AMI-related CS include acute myocarditis, arrhythmias, cardiac tamponade, traumatic cardiac injury, pulmonary thromboembolism (PTE), dynamic left ventricular (LV) outflow tract obstruction, myocardial depression in sepsis, and acute decompensated heart failure superimposed on chronic stable heart failure in conditions such as valvular heart disease, dilated cardiomyopathy, and constrictive pericarditis.1)4)5) CS is associated with a high rate of in-hospital mortality, with in-hospital mortality rate of AMI complicated by CS at 33–45%.6)7)8) Therefore, successful treatment of CS will result in an increased survival rate. To treat CS, various mechanical circulatory support (MCS) devices have recently been introduced, although medications such as inotropes, vasopressors, or diuretics are still priority treatment options. Intra-aortic balloon pump (IABP) has become the most widely used hemodynamic support device since its introduction in the 1960s.9) However, IABP was replaced by other advanced MCS devices during the recent decade. In Korea, the use of IABP decreased by >50% between 2012 and 2017. However, the use of extracorporeal membrane oxygenation (ECMO) has dramatically increased in the country (Figure 1). MEDICAL MANAGEMENT OF CARDIOGENIC SHOCK: The key principles of the medical management of CS are as follows: (i) maintenance of adequate blood pressure (BP) and organ perfusion, (ii) management of organ failure using ventilator and hemodialysis, and (iii) identification and resolution of the cause of cardiac failure. To stabilize a patient in CS, physicians attempt to optimize volume status and infuse inotropes and vasopressors to maintain adequate BP and organ perfusion. Although several inotropes, vasopressors, and combination methods have been used to treat CS, their selection is generally based on a physician's judgement and experience. Therefore, their use in a patient with CS is extremely diverse. Subgroup analysis in a multicenter, randomized trial showed that dopamine, compared to norepinephrine, was associated with an increased early mortality rate among patients with CS.10) The role of medical therapy may be confined to the stabilization of the patient until the cause is resolved. Current inotropes for CS may complicate AHF by increasing myocardial oxygen demand and consumption, and vasoconstrictors may impair microcirculation or tissue perfusion of the heart and other major organs11); therefore, their use should be restricted to the shortest possible duration and the lowest possible dose.5) The ESC recommended that short-term intravenous infusion of inotropes may be considered in patients with CS (class IIb recommendation).1) Failure of maximal medical therapy is no longer a justifiable endpoint, given the array of available mechanical circulatory supports.12)13)14)15) TIMING OF MECHANICAL SUPPORT: STABILIZATION PRIORITY STRATEGY: Although the most studied disease among the causes of CS is AMI, there are only a few randomized trials for AMI complicated by CS in which better management options were investigated. In 1999, Hochman et al. (the SHOCK investigators) found that early revascularization did not improve 30-day survival.16) In 2012, Thiele et al.17) (IABP-SHOCK II investigators) reported that compared with medical management, IABP did not improve 30-day survival. Thiele et al.18) (CULPRIT-SHOCK investigators) revealed that the 30-day survival of culprit-only revascularization is superior to that of complete revascularization in patients with CS caused by AMI. Because the most common cause of 30-day death in these patients is multi-organ failure due to profound shock or cardiac arrest, maintaining adequate hemodynamic support during the initial stage may be crucial. Therefore, short-term MCS other than IABP has been recommended in medically refractory CS to unload the failing ventricle and maintain sufficient end-organ perfusion; however, the level of recommendations and evidence in the guidelines of major cardiology societies is low.1)19)20)21) Esposito et al.22) introduced the term “door-to-support time” and recommended that this time should be reduced. In fact, the most important factor in the delay of “door-to-support time” is thought to be the decision-making time taken to implement MCS rather than the MCS procedure time. The SHOCK and CULPRIT-SHOCK trials showed that the timing and completeness of revascularization of a coronary artery may not be the critical factor for early survival of patients with AMI and CS. The IABP-SHOCK II trial revealed that IABP may be insufficient to support such unstable patients, i.e., 45% of patients had cardiopulmonary resuscitation (CPR) before randomization. These facts support the recommendation of strong MCS followed by cause correction. Cardiac output, BP, and systemic vascular resistance have a compound effect on systemic circulation. Esposito et al.22) described heart failure that is not complicated by CS as a “hemodynamic” problem and heart failure complicated by CS as a “hemometabolic” problem. If hemodynamic derangements persist in patients with AHF, a potentially reversible hemodynamic problem transitions to a more complex “hemometabolic” problem that may not respond to hemodynamic support alone. We believe that the decision-making time is usually extended by administering inotropes or vasopressors to the maximum dosage. However, as mentioned previously, most medications for CS increase myocardial oxygen demand and impair microcirculation. Basir et al.23) described that early implantation of an MCS device before the use of inotropes and vasopressors was associated with increased survival (p=0.05). Survival was 68%, 46%, 35%, 35%, and 26% for patients requiring 0, 1, 2, 3, and ≥4 inotropes before MCS support, respectively (p<0.001). Therefore, we believe that the prognosis would be optimal if the decision regarding MCS implementation occurs as soon as AHF becomes complicated by CS or when the end-organ perfusion begins to decrease. Indicators of end-organ malperfusion are blood lactate level, mixed venous saturation, and urine output.22)24) Certainly, early MCS implementation before the patient requires high-dose medications is theoretically attractive but requires further study in controlled trials. Several recent studies have reported improved survival with early initiation of MCS before percutaneous coronary revascularization in patients with CS complicating AMI.23)25)26)27)28) In addition, comprehensive consideration should be given to the risk of complications of MCS, possibility of MCS weaning (bridge to recovery), and intention of the patient and family members to bridge the long-term MCS (bridge to bridge or destination therapy) or heart transplantation (bridge to transplant). In addition, MCS devices can be used as a “bridge to decision or diagnosis.”29) MECHANISM: This device uses the counterpulsation of a balloon in the descending thoracic aorta to decrease the LV afterload and increase coronary perfusion. HEMODYNAMIC EFFECTS: Inflation occurs at the onset of diastole, just after the dicrotic notch on the aortic pressure waveform, which increases the diastolic pressure of the aorta. There is a resulting increase in systemic perfusion, including coronary perfusion. Balloon deflation is timed to occur at the end of diastole or immediately prior to systole and reduces the LV afterload leading to improved cardiac output (Figure 2A).30) COMPONENTS AND CATHETERIZATION: The balloon catheter is inserted through the femoral artery and positioned in the proximal descending thoracic aorta. The system comprises a dual-lumen 7.5-F to 8.0-F catheter with a polyethylene balloon and the control console. The inner catheter lumen accepts the guidewire during placement and transduces the aortic pressure for monitoring. The gas lumen serves as the conduit for the rapid exchange of helium in and out of the balloon.30) The proper position of the distal tip is 1–2 cm distal to the left subclavian artery (the second to third intercostal space or 2 cm above the carina).31) Balloon obstruction of the left subclavian artery and visceral arteries, such as the superior mesenteric artery and renal arteries, should be avoided, and balloon positioning should be performed under fluoroscopy to avoid these potential complications. If the placement cannot be performed under fluoroscopy, immediate verification via a plain chest film is warranted. Serial monitoring of the left radial pulse and urine output can indicate IABP malposition. MANAGEMENT: There are 2 modes for setting the balloon trigger, namely, electrocardiogram and pressure. Scenarios including early or late timing of inflation or deflation are undesirable. Early inflation or late deflation leads to increased afterload and increased myocardial oxygen demand. Late inflation or early deflation does not sufficiently increase coronary perfusion. COMPLICATIONS: Complications include balloon rupture, bowel ischemia, and stroke.32) CONTRAINDICATIONS: Contraindications to IABP placement include severe aortic insufficiency, aortic disease, and peripheral vascular disease.32) Unlike other MCS devices, an IABP needs synchronization with the ventricular cycle. Therefore, the implementation is restricted in case of cardiac arrest or severe tachy- or brady-arrhythmia. CLINICAL OUTCOMES AND INDICATIONS: In 2013, the American College of Cardiology and the American Heart Association recommended the use of an IABP in patients with CS after ST-elevation myocardial infarction who do not quickly stabilize with pharmacological therapy (IIa, B)19); however, contrary results have been reported. IABP does not provide any survival benefit to high-risk patients with CS because it only provides modest hemodynamic support (0.5–0.8 L/min).33) In the intra-aortic balloon counterpulsation in patients with AMI complicated by CS (IABP-SHOCK I) trial, the use of IABP for hemodynamic stability was not beneficial.34) In addition, the IABP-SHOCK II trial concluded no mortality benefit of IABP compared to that of medical therapy.35) The results of the IABP-SHOCK II trial influenced the 2014 ESC guidelines with a further downgrading of IABP with a new class III (harm) recommendation for routine use in CS.36) The ESC still does not recommend the routine use of IABP in patients with CS (III, B).1)20)21) The use of IABP support for patients with AMI-associated mechanical complication of acute mitral regurgitation or ventricular septal rupture remains a class II indication in the 2017 ESC guidelines.20) Some doctors do not include IABP in MCS devices because IABP is limited in its effect to reducing cardiac afterload and increasing coronary perfusion rather than increasing the systemic circulation. We suggest that the indication for the use of IABP should be limited to only those patients with AHF not complicated by CS. RIGHT ATRIAL-TO-AORTA ACCESS DEVICE: EXTRACORPOREAL MEMBRANE OXYGENATION: All ECMO cases described in this article used the venoarterial (VA) type of ECMO. MECHANISM: ECMO involves the withdrawal of deoxygenated blood from the venous system through a drainage cannula, pumping the blood through a membrane lung, and returning the blood to the arterial circulation through a return cannula. HEMODYNAMIC EFFECTS: ECMO does not directly decompress the left ventricle. As blood is pulled from the venous system, there is a decrease in the right ventricular (RV) preload and consequently the LV preload, which can reduce wall tension and workload. In contrast, as blood returns to the arterial system in a retrograde manner, there is an increase in the LV afterload and workload.24) Left ventricular end-diastolic pressure (LVEDP) may increase as the ECMO flow increases.37) Left ventricular end-diastolic volume (LVEDV) may increase or decrease depending on the balance between the decreased LV preload and remnant native LV contractility to overcome the increased LV afterload. Without satisfactory ejection, blood will accumulate under pressure until it eventually equalizes with systemic arterial pressure (Figure 2B). COMPONENTS AND CANNULATION: The basic ECMO circuit includes cannulae for drainage and return, tubing, a pump, and a membrane lung. The cannulae can be placed either centrally or peripherally, but most ECMO is cannulated peripherally. Peripheral cannulation entails drainage of the venous blood from the right atrium (RA) through a cannula that exits through the femoral vein. The arterial cannula is a short cannula inserted into the femoral artery with the tip in the common iliac artery. It is relatively easy to place and can be inserted without fluoroscopy, making it an ideal device to support patients receiving ongoing CPR.4) MANAGEMENT: A major advantage of ECMO is the relative ease of implementation; however, some disadvantages include the need for specialized perfusion expertise and skillful nursing care.38) The reason for specialized skill and expertise is because during ECMO, it is necessary to monitor the recovery of adequate systemic perfusion while maintaining sufficient ECMO flow and adequate BP. COMPLICATIONS: Complications include bleeding, local vascular injury, limb ischemia, thromboembolism, LV distension, pulmonary edema, and air embolism.24)49)50) CONTRAINDICATIONS: Contraindications for the placement of the ECMO circuit include severe aortic regurgitation (AR) and peripheral arterial occlusive disease.32) CLINICAL OUTCOMES AND INDICATIONS: While the use of ECMO has increased in the USA from 2004–2014, the outcomes remain poor with an in-hospital mortality of 47% in 2014.51) The usage trend of ECMO in Korea has also increased, and the in-hospital mortality was 63.4% between 2009 and 2014 (Figure 1B).52) Muller et al.53) assessed the clinical and quality of life outcomes of ECMO in 138 AMI patients with CS who did not require CPR. Based on multivariable logistic-regression analyses of data of this cohort of patients, the ENCOURAGE score to predict mortality in ECMO-treated AMI patients was calculated using the following pre-ECMO parameters: age >60 years, female sex, BMI >25 kg/m2, Glasgow coma score <6, creatinine >150 μmol/L, lactate <2, 2–8 or >8 mmol/L, and prothrombin activity <50%. Six months after ECMO, the probabilities of survival were 80%, 58%, 25%, 20%, and 7% for ENCOURAGE score classes 0–12, 13–18, 19–22, 23–27, and at least 28, respectively.53) Combes et al.54) analyzed 65 patients receiving ECMO (72% of patients were cannulated percutaneously) for CS secondary to dilated cardiomyopathy, fulminant myocarditis, post-cardiotomy, post-transplantation, and other conditions; their in-hospital mortality was 58%. Because it is less expensive than other devices, allows rapid improvement in oxygenation, and is the only short-term MCS suitable for patients with severe biventricular failure or cardiac arrest, ECMO has emerged as the first-line support system, with a growing number of accepted indications.55) The typical indications for the use of ECMO include the following: MECHANISM: The TandemHeart device (CardiacAssist Inc., Pittsburgh, PA, USA) is a support device that delivers blood from the LA to the arterial system using an extracorporeal centrifugal pump. The TandemHeart works in tandem with the left ventricle. Outflow from the TandemHeart device and left ventricle run in tandem, but toward each other.30) HEMODYNAMIC EFFECT: During the TandemHeart support, the LV preload decreases and the LV afterload increases (Figure 2C). COMPONENTS AND CANNULATION: The TandemHeart system is composed of a drainage cannula, return cannula, centrifugal pump, and control console. The drainage cannula is placed from the femoral vein into the LA, whereas the return cannula is inserted retrogradely into the femoral artery. The placement of the drainage cannula requires a transseptal puncture. For an experienced operator, performing a transseptal puncture is a relatively safe procedure. MANAGEMENT: The flow provided by the TandemHeart device is dependent on several factors, including the systemic and pulmonary resistance, cannula size and position, and fluid balance. Adequate RV function is required to maintain the left atrial volume.30) Continued hemodynamic monitoring, as well as the utility of a Swan-Ganz catheter, help to assess the total cardiac output and filling pressures. Vibration in the system's tubing may signal inadequate filling of the LA and should trigger an evaluation of the root cause, which may include hypovolemia, pulmonary hypertension, cardiac tamponade, bleeding, RV failure, or arrhythmias. Kinks in the tubing, cannula migration, and thrombus in the circuit should also be assessed. Rarely, the left atrial cannula can migrate into the RA, causing hemodynamic collapse and profound hypoxia.4) If the pulmonary edema and hypoxemia complicating AHF is severe, it is possible to insert a membrane lung into the TandemHeart circuit, making it similar to ECMO.77) Adequate anticoagulation is important because the drainage cannula could form thrombi in the LA. The manufacturer of the TandemHeart device recommends using a heparinized purge solution. Unlike the Impella devices, the TandemHeart device purge solution is released at a fixed rate.78) COMPLICATIONS: Complications include bleeding, local vascular injury, limb ischemia, air embolism, thromboembolism, device dislodgement, arrhythmias, cardiac wall perforation, aortic root puncture, pericardial effusion or cardiac tamponade, stroke, and residual atrial septal defects.4)30)38)79) CONTRAINDICATIONS: Contraindications include any condition that prohibits anticoagulation. In addition, the presence of a left atrial thrombus, ventricular septal defect, or aortic insufficiency precludes placement.4)38)79) CLINICAL OUTCOMES AND INDICATIONS: In an analysis of 117 patients with severe refractory CS, Kar et al.80) observed significant improvements in the cardiac index, systolic BP, mixed venous oxygen saturation, and urine output with the use of TandemHeart support. Thirty-day and 6-month mortality rates were 40.2% and 45.3%, respectively, even though 47.9% patients underwent CPR immediately before or at the time of implantation. Burkhoff et al.81) randomized 33 patients within 24 hours of developing CS to treatment with an IABP or TandemHeart device. Compared with the IABP group, the TandemHeart group showed a greater increase in the cardiac index and mean arterial BP and a decrease in the pulmonary capillary wedge pressure; however, no difference in severe adverse events or 30-day mortality was observed between the 2 groups. Alli et al.82) reported a series of 54 patients who received support of TandemHeart device while undergoing high-risk percutaneous coronary intervention (PCI). The procedural success was high at 97%, and the 6-month survival was 87%. However, major vascular complications occurred in 13% of patients. LEFT VENTRICULAR-TO-AORTA ACCESS DEVICE: THE IMPELLA DEVICE: The Impella device (Abiomed Inc., Danvers, MA, USA) and HeartMate Percutaneous Heart Pump (Thoratec Corporation, Abbott Laboratories, Pleasanton, CA, USA) are current percutaneous ventricular assist devices (VADs). The Impella device that is described in this article is a left VAD, and a description of the Impella RP, which is a right VAD, is provided later. MECHANISM: The Impella is a percutaneous VAD that uses a microaxial pump to move blood continuously from the left ventricle to the ascending aorta. The device utilizes a constant axial flow, does not need electronic synchronization, and permits steady output regardless of arrhythmias.83) HEMODYNAMIC EFFECTS: There are 2 primary effects of the Impella device, including (i) unloading of the ventricle (lower end-diastolic volume and pressure) and (ii) an increase in forward flow (higher mean arterial pressure) (Figure 2D). There is a reduction in LVEDP and LVEDV that translates into a decrease in the wall tension and myocardial oxygen demand.38)79) The hemodynamic effect of the Impella device is similar to the effect of an IABP, however, the amount of circulatory support is considerably higher than that of IABP. The Impella device effectively empties the left ventricle, lowers the end-diastolic pressure, and increases the peak coronary blood flow, resulting in an increased supply and decreased demand of myocardial oxygen.84) COMPONENTS AND CATHETERIZATION OR CANNULATION: There are three types of Impella devices that are available to provide increasing levels of LV support, namely, the Impella 2.5 (2.5 L/min; 12Fr motor pump), the Impella CP (approximately 3.5 L/min; 14Fr motor pump), and the Impella 5.0 (5.0 L/min; 21Fr motor pump). All three devices have been approved in the United States to provide hemodynamic support for up to 6 hours. The Impella support system is comprised of 3 major components, namely, (i) a catheter, (ii) a purge system, and (iii) the automated controller. An impeller and its adjacent motor are located near the outlet area in the ascending aorta. As it rotates, negative pressure draws ventricular blood into the inlet area and through the cannula. The rotation speed of the impeller has nine settings (P0 to P8). To protect the motor, the purge fluid (5% dextrose with heparin) forms a hydraulic pressure shield that prevents blood from migrating past the impeller and into the motor housing. Sizes of the catheters, motor pumps, and introducers and access methods for each type are presented in Table 1.85) The motor pump of the Impella 5.0 device (21F) requires an end-to-side anastomosis of the 10-mm Dacron graft to the femoral or axillary artery.30) COMPLICATIONS: Abaunza et al.86) reported acute limb ischemia in up to 13% of patients after they had inserted the Impella 2.5 or CP, which was related to emergency procedures. Other complications included bleeding at the vascular access site, which required transfusion in 24.2% of cases and vascular surgery in 4.2% of patients. Hemolysis resulting in blood transfusion was reported in 7.5% of patients. In 1.7% of patients, pericardial drainage was necessary because of cardiac tamponade. Device malfunction that necessitated the removal of the device occurred in 2.5% of patients during long-term support.87) CONTRAINDICATIONS: The presence of a mechanical aortic valve or LV mural thrombi precludes the use of the Impella device, as does significant aortic valve stenosis. These clinical parameters do not prohibit the use of the TandemHeart device or ECMO.30) The use of the Impella device is also contraindicated in patients with moderate-to-severe AR, ventricular septal defect, and severe peripheral vascular disease.4) CLINICAL OUTCOMES AND INDICATIONS: An evidence-based efficacy study conducted in patients with CS showed that the Impella device can improve cardiac output and the cardiac index; however, no significant reduction in mortality was observed.88)89) Seyfarth et al.90) randomly allocated 25 patients with AMI and CS to receive percutaneous support with either an IABP or the Impella 2.5 device and showed that the cardiac index after 30 min of support was greater with the use of the Impella. However, the overall 30-day mortality was 46% in both groups. The PROTECT 2 trial was a randomized clinical trial that enrolled 452 patients who underwent high-risk PCI with the support of an IABP or the Impella 2.5 device. There was an increased number of 30-day major adverse events in the IABP group. At 90 days, a strong trend toward decreased major adverse events was observed in the Impella 2.5-supported patients compared with that in the IABP group. This trial was terminated at 69% of the planned enrollment for futility.91) MECHANICAL RIGHT VENTRICULAR ASSIST DEVICES: For patients with RV failure complicated by CS, percutaneous RV MCS may be necessary. Diagnosing acute RV failure remains a major clinical challenge. Invasive hemodynamic measures can be obtained using a pulmonary artery (PA) catheter and are predictive of RV failure. The simplest approach to quantify RV dysfunction is to measure the ratio of RA pressure to pulmonary capillary wedge pressure.92) In the setting of AMI, Lopez-Sendon et al.93) identified that a ratio >0.86 was associated with pathological evidence of RV infarction at necropsy. Korabathina et al.94) reported the clinical utility of the PA pulsatility index (<1.0) as a measure of RV failure in the setting of AMI. For a direct RV bypass, the Impella RP or TandemHeart using a Protek-Duo catheter (CardiacAssist Inc.) is appropriate. For an indirect RV bypass, ECMO support is suitable.4) THE IMPELLA RP DEVICE: The Impella RP device is implanted into the femoral vein for inflow through the inferior vena cava to the outlet area in the PA. The Impella RP device is Food and Drug Administration approved for patients who develop acute right heart failure or decompensation after left VAD implantation, myocardial infarction, heart transplantation, or open-heart surgery (Figure 2E). To achieve biventricular support, Pappalardo et al.95) reported the first case of biventricular support using 2 Impella pumps, that is, a combination of an Impella CP with an Impella RP system, for acute biventricular failure due to suspected acute myocarditis. This concept extends the possibilities of different unloading strategies for patients with biventricular failure in whom oxygenation is not a major issue.96) For example, for right heart failure due to massive PTE, even if the Impella CP and RP devices are supported at the same time, the blood of the main or proximal PA cannot be transmitted to the pulmonary veins through the thromboemboli in the pulmonary arteries. To date, the ECMO device is the only MCS device for massive PTE complicated by CS. SELECTION OF MECHANICAL CIRCULATORY SUPPORT: Multiple factors must be considered when choosing MCS, including the hemodynamic condition of the patient; hemodynamic effects; indications or limitations of the devices; technical considerations, including ease and rapidity of insertion; and the ultimate goals of support.15)38)97) Some considerations that we believe should be taken into account are as follows: The algorithm for MCS selection is presented in Figure 5, reflecting the various factors above. FUTURE DIRECTIONS: Technological developments are ongoing and include simpler and faster cannulation or catheterization, less anticoagulation, more ventricular unloading effect, and easier management with less complications. There are multiple open questions regarding MCS devices, as reflected by the high number of recommendations with a level of evidence C (limited data or expert opinion) in current guidelines.19)36)98) Moreover, randomized clinical trial
PMID:36263076 | 1 | 36263076 | PMC9536734 | 10.36628/ijhf.2019.0015 | fetched |  | s on CS are difficult to perform, and only a few randomized clinical trials came to clinical outcomes with completion of the required patient number.16)17) This should be the motivation for future randomized clinical trials involving MCS devices. CONCLUSIONS: MCS for patients with AHF complicated by CS has been increasing in popularity. Proper device selection with appropriate timing of application is important. Considering that various factors are involved in the complexity of decision making, the hospital should have its own standardized approach to patients with AHF complicated by CS.
PMID:38533653 | 0 | 38533653 | PMC11097944 | 10.1161/CIRCINTERVENTIONS.123.013898 | fetched | {"title": "TABERNACL: Temporary Hemodynamic Stabilization In Vivo.", "authors": ["Halaby RN", "Bruce CG", "Yildirim DK", "Uzun D", "Rogers T", "Khan JM", "Jaimes AE", "Grant LP", "Babaliaros VC", "Greenbaum AB", "Lederman RJ"], "journal_abbrev": "Circ Cardiovasc Interv", "year": 2024, "volume": "17", "issue": "5", "pages": "e013898", "doi": "10.1161/CIRCINTERVENTIONS.123.013898", "pmid": "38533653", "pmcid": "PMC11097944", "pub_types": ["Journal Article", "Research Support, N.I.H., Intramural"]} | TABERNACL: Temporary Hemodynamic Stabilization In Vivo. BACKGROUND: Acute aortic regurgitation is life-threatening with few nonsurgical options for immediate stabilization. We propose Trans-Aortic Balloon to Ease Regurgitation Applying Counter-Pulsation (TABERNACL), a simple, on-table temporary valve using commercially available equipment to temporize acute severe aortic regurgitation. METHODS: We hypothesize that an appropriately sized commercial balloon dilatation catheter-straddling the aortic annulus and connected to a counterpulsation console-can serve as a temporizing valve to restore hemodynamic stability in acute aortic regurgitation. We performed benchtop testing of valvuloplasty, angioplasty, and sizing balloons as counterpulsation balloons. TABERNACL was assessed in vivo in a porcine model of acute aortic regurgitation (n=8). We also tested a static undersized, continuously inflated transvalvular balloon as a spacer intended physically to obstruct the regurgitant orifice. RESULTS: Benchtop testing identified that Tyshak II and PTS sizing (NuMed Braun) balloon catheters performed adequately as temporary valves (ie, complete inflation and deflation with each cycle) and resisted fatigue, in contrast to others. When TABERNACL was used in the acute severe regurgitation animals, there was immediate hemodynamic improvement, with a significant 35% increase in diastolic aortic pressure by 16 mm Hg ([95% CI, 7-25] P=0.0056), 34% reduction in left ventricular end-diastolic pressure by -7 mm Hg ([95% CI, -10 to -5] P=0.0006), improvement in the aortic diastolic index by 0.28 ([95% CI, 0.18-0.39] P=0.0009), and reversal of electrocardiographic myocardial ischemia. As an alternative, static balloon inflation across the aortic valve stabilized regurgitation hemodynamics at the expense of a new aortic gradient and caused excessive ectopy from balloon movement in the left ventricular outflow tract. CONCLUSIONS: TABERNACL improves hemodynamics and reduces coronary ischemia by electrocardiography in animals with acute severe aortic regurgitation. TABERNACL valves obstruct the diastolic regurgitant orifice without systolic obstruction. This may prove a lifesaving bridge to definitive valve replacement therapy.

## _meta (hidden)
schema | claim-check/1
title | IABP temporización (HUD UCI)